| Confidentiality Level |  | BL - Restricted for internal use | TC ID / Revision |  | 00200314/C |
| --- | --- | --- | --- | --- | --- |
| Document Status |  | Document Released | Document No. |  | N/A |
| WBS code |  | 3.2. - L2 system |  |  |  |
| PBS code |  | RA1.L2.L2_1.MIR.DG.SPM.2 |  |  |  |
| Project branch |  | Engineering & Scientific documents (E&S) |  |  |  |
| Document Type |  | Specification (SP) |  |  |  |
| [RSD product category A] NIR spectrometer TP18_799 Keywords N/a |  |  |  |  |  |
|  | Position |  |  | Name |  |
| Responsible person | Senior Researcher / Laser 1 Laser Physicist |  |  | Jonathan Tyler Green Radek Šmíd |  |
| Prepared by | Laser Physicist |  |  | Radek Šmíd |  |

| RSS TC ID/revision | RSS - Date of Creation | RSS - Date of Last Modification | Systems Engineer |
| --- | --- | --- | --- |
| 015296/A.001 | 01-Nov-2018 10:56 | 01-Nov-2018 11:13 | Aleksei Kuzmenko |
| 015296/A.002 | 12-Nov-2018 16:57 | 12-Nov-2018 16:57 | Aleksei Kuzmenko |
| 015296/A.003 | 21-Nov-2018 08:33 | 21-Nov-2018 08:33 | Pavel Tůma |

| Reviewed By |  |  |  |
| --- | --- | --- | --- |
| Name (Reviewer) | Position | Date | Signature |
| Bedřich Rus | Scientific Coordinator of Laser Technology (RP1) | NOTICE (RSD product category A) |  |
| Daniel Kramer | Team Leader Scientific OD | NOTICE (RSD product category A) |  |
| Ladislav Půst | Manager installation of technology | NOTICE (RSD product category A) |  |
| Martin Laub | Chief Engineer | NOTICE (RSD product category A) |  |
| Pavel Bakule | Deputy RP1 Leader | NOTICE (RSD product category A) |  |
| Roman Kuřátko | Facility Manager | NOTICE (RSD product category A) |  |
| Veronika Olšovcová | Safety Coordinator | NOTICE (RSD product category A) |  |
| Viktor Fedosov | SE & Planning group leader; Quality Manager | NOTICE (RSD product category A) |  |

| Approved by |  |  |  |
| --- | --- | --- | --- |
| Name (Approver) | Position | Date | Signature |
| Bedřich Rus | Scientific Coordinator of Laser Technology (RP1) |  |  |

| Revision History / Change Log |  |  |  |  |
| --- | --- | --- | --- | --- |
| Change No. | Made by | Date | Change description, Pages, Chapters | TC rev. |
| 1 | R. Šmíd | 19.10.2018 | RSD draft creation | A |
| 2 | J.T. Green, A.A. Kuzmenko | 12.11.2018 | RSD update, version for internal review | B |
| 3 | P. Tůma | 21.11.2018 | RSD update, final version for approval | C |

Content

1. Introduction	4

1.1. Purpose	4

1.2. Scope	4

1.3. Terms, Definitions and Abbreviations	4

1.4. Reference documents	4

1.5. References to standards	5

2. General Functional and Performance requirements	5

3. Delivery requirements	6

4. Safety Requirements	6

5. Quality Requirements	7

5.1. General Quality Requirements	7

# Introduction

## Purpose

This Requirements Specification Document (RSD) lists the technical requirements and constraints for the selection of Near-infrared spectrometer used in DUHA projects within the ADONIS grant.

## Scope

This RSD contains all of the technical requirements: functional, performance, delivery, safety and quality requirements for the following product (tender number: TP18_799): NIR spectrometer (further “Spectrometer”).

The Spectrometer is considered to be the standalone technology and will be all placed in the L1 Oscillator room. This Spectrometer is registered in the PBS database under the following PBS code: RA1.L2.L2_1.MIR.DG.SPM.2.

The Spectrometer is a product Category A according to the ELI Beamlines RSD categories. The category A is an Off-the-shelf Product without the necessity of modifications and necessity to be subjected to a verification programme (review of design, inspection and testing) for ELI applications by the actual project specifications. All verification activities performed by a supplier shall be executed in accordance with the supplier’s plan of outgoing inspection and tests. The verification of all specified parameters listed in this RSD shall be undertaken by the supplier before delivery to the ELI Beamlines facility. Furthermore, the Spectrometer may be subject to testing and verification upon delivery to the ELI Beamlines facility by qualified personnel. All non-conformances (if any) must be addressed by the supplier in a timely manner.

## Terms, Definitions and Abbreviations

For the purpose of this document, the following abbreviations apply:

| Abbreviation | Meaning |
| --- | --- |
| A/D | Analog-to-Digital |
| API | Application Programmers Interface |
| CA | Contracting Authority (Institute of Physics AV CR, v. v. i.) |
| ELI | Extreme Light Infrastructure |
| L1 | Laser 1 |
| NIR | Near-Infrared |
| PBS | Product Breakdown Structure |
| RSD | Requirements Specification Document |
| TTL | Transistor–Transistor Logic |
| USB | Universal Serial Bus |

## Reference documents

| Number of document | Title of Document/ File |
| --- | --- |
| RD-01 | 00163567-B_1.2_Q_M_Guide_for_Instructions_for_Use_Ver-9_EN_fully_signed.pdf |

## References to standards

If this document includes references to standards or standardized/ standardizing technical documents the CA allows/permits also another equal solution to be offered. If a supplier offers another equal solution the CA shall not reject its bid, once the supplier by appropriate means in the bid proves that the offered supplies, services or works meet in an equivalent manner the requirements including references to standards or technical documents.

# General Functional and Performance requirements

REQ-025175/A

The Spectrometer shall meet all the requirements defined in Table 1 below.

| Description of parameter | Value |
| --- | --- |
| Wavelength range | 900 – 2500 nm |
| Spectral resolution | <5 nm |
| Fibre Input | SMA 905 |
| Slit dimension | 10 m |
| Detector type | Linear InGaAs |
| Signal-to-Noise ratio | 10000:1 @ 100 ms |
| Integration time minimum range | 1 ms – 100 ms |
| Internal Moving parts | NO |
| External trigger | YES – compatible with TTL logic input |
| A/D resolution | 16 bit |
| Stabilization of thermal drifts of the spectra | YES |
| 2nd order filter | YES |
| Operating temperature | 10-35 °C |
| Connectivity | USB2 |
| Software requirements | Must include LabVIEW drivers and LabVIEW API in accordance with the LabVIEW Instrument Driver Guidelines with proven compatibility with LabVIEW Real-Time operating systems (see e.g., ni.com/idnet) |
| Electric Power consumption | USB power input 5 V |
| Others | The calibration data for wavelength and sensitivity (see REQ-025413/A) |

Table 1: Functional and performance parameters of the Spectrometer.

# Delivery requirements

REQ-025176/A

The transportation to the final destination of the Spectrometer shall be conducted by the Supplier.

REQ-025177/A

The Spectrometer and its components shall be suitable for use in cleanrooms of class 7 according to ČSN EN ISO 14644 (equivalent to EN ISO 14644). The Spectrometer and its components shall be delivered in a protective package preventing damage and contamination and in a minimum of two plies of antistatic clean packaging.

# Safety Requirements

REQ-025179/A

The Supplier shall supply a Declaration of Conformity or any other equivalent document legally recognized and accepted in the Czech Republic for each product type if the appropriate legislation determines the Supplier's obligation to have a Declaration of Conformity (or the equivalent document) for the purposes of a Product sale in the Czech Republic to fulfil the requirements of 2001/95/EC directive or applicable Czech law.

# Quality Requirements

## General Quality Requirements

REQ-025180/A

The Supplier shall provide Instructions for use (Product User Manual) as part of the delivered Product. The Instructions for use shall be written in accordance with standard ČSN EN 82079-1 (or equivalent, e.g. EN 82079-1) and shall include the instructions and descriptions regarding the following:

- transport, handling and storage;

- installation, cleaning and calibration (see REQ-025413/A);

- safe operation and maintenance procedures;

- user manual for the software or libraries and for communication protocols.

NOTE: As an alternative to standard ČSN EN 82079-1 (or equivalent, e.g. EN 82079-1) an internal ELI “Instructions for use” methodology can be used (see RD-01; chapter 1.4) which will be provided to the Supplier upon request.

REQ-025413/A

The Supplier shall provide a Calibration Certificate giving details of the Spectrometer amplitude and wavelength calibration, e.g. calibrations method and results.

NOTE: Recommended calibration interval for the Spectrometer shall be defined and documented in the Calibration Certificate or Instructions for use (see REQ-025180/A).

REQ-025181/A

The Supplier shall provide information on the execution of outgoing check of the Product. At least this information shall comprise a declaration about the execution of outgoing check and declaration of conformity with technical requirements defined by the product RSD and completeness of the Product.

REQ-025182/A

The Supplier shall establish and maintain a nonconformity control system compatible with ČSN EN ISO 9001 (or equivalent, e.g. EN ISO 9001).