- Technical Specifications

Tender: „Mechanical helium cryogenic system“

Optical Cryostat with highly accurate temperature control

## Manufacturer: Montana Instruments

Model:  Customized system based onstandard system Cryostation C2

Technical Specifications

- Sample temperature range: < 4,5 K – 350 K

- Sample temperature stability at the detector mount <50 µK

- Vibration stability at the base of the sample chamber < 5 nm (peak-to-peak)

- Desk-top solution of the cryostat. Version for metric optical tables

- Sample Chamber dimensions:

  - diameter 53 mm inside the cold shield

  - height 100 mm inside the cold shield

- Cooled chamber includes two temperature sensors Cernox

- Cold head includes two temperature sensors

- Air-cooled compressor with variable speed for lowest noise. Loudness < 65 dB

- Mechanical cryo-cooler, air-cooled

- Cooling time from room temperature down to 4,5 K in 4 hours

- 4 vacuum feedthroughs and 4 radiation windows

- 28 electrical feedthroughs + 1 ground via mini pin connectors

- 16 wires vie Fisher connectors, thermally anchored at the 35 K and 4 K stages to prevent parasitic heat leaks.

- 2 compression feedthroughs for 125 µm or 250 µm fiber

- 10 m Helium hoses

- Radiations shields: 35K and 4 K inner isothermal shield for sample chamber

- Helium pressure lines galvanically isolated

- Control unit

- User interface computer with software for automated operation (licence for one computer)