# Product specification

Universal Oven Um

UF110mplus

# The heating oven Um is a Class I medical device.

This universal oven made of high-quality hygienic, easy to clean stainless steel leaves nothing to be desired in terms of ventilation technology, control technology, overtemperature protection and perfectly tuned heating technology.

## Temperature

Working temperature range	at least 5 (UN/UNplus/UNm/UNmplus) or 10 (UF/UFplus/UFm/UFmplus) above ambient temperature to +300 °C

Setting accuracy temperature

up to 99.9 °C: 0.1 / from 100 °C: 0.5

Setting temperature range	+20 to +300 °C

Temperature sensor	2 Pt100 sensors DIN Class A in 4-wire-circuit for mutual monitoring, taking over functions in case of an error

## Control technology

ControlCOCKPIT	TwinDISPLAY. Adaptive multifunctional digital PID-microprocessor controller with 2 high-definition TFT-colour displays.

Language setting	German, English, Spanish, French, Polish, Czech, Hungarian

Timer	Digital backwards counter with target time setting, adjustable from 1 minute to 99 days

Function HeatBALANCE	adapting the distribution of the heating performance of the upper and lower heating circuit from -50 % to +50 %

Function SetpointWAIT	the process time does not start until the set temperature is reached

Calibration	three freely selectable temperature values

adjustable parameters	temperature (Celsius or Fahrenheit), fan speed, air flap position, programme time, time zones, summertime/wintertime

## Ventilation

Fan	forced air circulation by quite air turbine, adjustable in 10 % steps for each segment individually

Fresh air admixture	adjustment of pre-heated fresh air admixture by air flap control in 10 % steps for each segment individually

Vent	vent connection with restrictor flap

## Communication

Documentation	programme stored in case of power failure

Programming	AtmoCONTROL software on a USB stick for programming, managing and transferring programmes via Ethernet interface or USB port

## Safety

Temperature control	mechanical temperature limiter TB, protection class 1 according to DIN 12880 to switch off the heating approx. 20°C above nominal temperature

Temperature control	overtemperature monitor TWW, protection class 3.1 or adjustable temperature limiter TWB, protection class 2, selectable on display

AutoSAFETY	additionally integrated over- and undertemperature monitor "ASF", automatically following the setpoint value at a preset tolerance range, alarm in case of over- or undertemperature, heating is switched off in case of overtemperature

Autodiagnostic system	for fault analysis

Alarm	visual and acoustic

## Standard equipment

Door	fully insulated stainless steel door with 2-point locking (compression door lock)

Internals	2 stainless steel grid(s), electropolished

Works calibration certificate	Calibration at +160°C

## Stainless steel interior

Dimensions	w(A) x h(B) x d(C): 560 x 480 x 400 mm (d less 39 mm for fan)

Interior	easy-to-clean interior,made of stainless steel, reinforced by deep drawn ribbing with integrated and protected large-area heating on four sides

Volume	108 l

Max. number of internals	5

Max. loading of chamber	175 kg

Max. loading per internal	20 kg

## Textured stainless steel casing

Dimensions	w(D) x h(E) x d(F): 745 x 864 x 584 mm (d +56mm door handle)

Housing	rear zinc-plated steel

## Electrical data

Voltage Electrical load

Voltage Electrical load

230 V, 50/60 Hz

approx. 2800 W

115 V, 50/60 Hz

approx. 1800 W

## Ambient conditions

Set Up	The distance between the wall and the rear of the appliance must be at least 15 cm. The clearance from the ceiling must not be less than 20 cm and the side clearance from walls or nearby appliances must not be less than 5 cm.

Altitude of installation	max. 2,000 m above sea level

Ambient temperature	+5 °C to +40 °C

Humidity rh	max. 80 %, non-condensing

Overvoltage category	II

Pollution degree	2

## Packing/shipping data

Transport information	The appliances must be transported upright

Customs tariff number	8419 8998

Country of origin	Federal Republic of Germany

WEEE-Reg.-No.	DE 66812464

Dimensions approx incl. carton

w x h x d: 830 x 1050 x 800 mm

Net weight	approx. 74 kg

Gross weight carton	approx. 99 kg

Standard units are safety-approved and bear the test marks