| Confidentiality Level |  | BL - Restricted for internal use | TC ID / Revision |  | 00148864/C |
| --- | --- | --- | --- | --- | --- |
| Document Status |  | Document Released | Document No. |  | N/A |
| WBS code |  | 5.3. - RP4 Applications in molecular, biomedical, and material sciences |  |  |  |
| PBS code |  | E.E1.HXRS.POSS - Positioning Systems for E1 Endstations 2 - X-rays |  |  |  |
| Project branch |  | Engineering & Scientific documents (E&S) |  |  |  |
| Document Type |  | Specification (SP) |  |  |  |
| [RSD product category B] Positioning systems for E1 Endstations 2 TP16_138 Keywords N/A |  |  |  |  |  |
|  | Position |  |  | Name |  |
| Responsible person | RP4 Team Leader |  |  | Jakob Andreasson |  |
| Prepared by | Researcher / Postdoc RP4 |  |  | Eva Klimešová |  |

| RSS TC ID/revision | RSS - Date of Creation | RSS - Date of Last Modification | Systems Engineer |
| --- | --- | --- | --- |
| 011973/A.000 | 11.04.2017 01:20 | 13.04.2017 21:30 | Marek Malý |
| 011973/A.001 | 13.04.2017 21:32 | 13.04.2017 21:33 | Marek Malý |
| 011973/A.002 | 24.04.2017 15:02 | 24.04.2017 15:04 | Marek Malý |
| 011973/A.003 | 02.05.2017 12:02 | 02.05.2017 12:03 | Aleksei Kuzmenko |

| Reviewed By |  |  |  |
| --- | --- | --- | --- |
| Name (Reviewer) | Position | Date | Signature |
| Jakob Andreasson | Group Leader RP4 | NOTICE (RSD product category B) |  |
| Christopher Brooks | Researcher RP4 | NOTICE (RSD product category B) |  |
| Martin Laub | Team Leader of Mechanics | NOTICE (RSD product category B) |  |
| Alice Hamalová | Clean Room Specialist | NOTICE (RSD product category B) |  |
| Ladislav Půst | Manager installation of technology | NOTICE (RSD product category B) |  |
| Luboš Nims | Head of Electrical engineering | NOTICE (RSD product category B) |  |
| Tomáš Laštovička | Group Leader BIS and C&C | NOTICE (RSD product category B) |  |
| Veronika Olšovcová | Safety coordinator | NOTICE (RSD product category B) |  |
| Viktor Fedosov | SE & Planning group leader; Quality Manager | NOTICE (RSD product category B) |  |

| Approved by |  |  |  |
| --- | --- | --- | --- |
| Name (Approver) | Position | Date | Signature |
| Georg Korn | Science and Technology Manager, Scientific coordinator of RP2-6 |  |  |

| Revision History / Change Log |  |  |  |  |
| --- | --- | --- | --- | --- |
| Change No. | Made by | Date | Change description, Pages, Chapters | TC rev. |
| 1 | E.Klimešová | 28.03.2017 | RSD draft creation | A |
| 2 | M.Malý | 13.04.2017 | RSD update, addition of standard chapters | A |
| 3 | E.Klimešová | 24.04.2017 | RSD update, version for internal review | B |
| 4 | E.Klimešová, A.Kuzmenko | 28.04.2017 | RSD update, final version for approval | C |

Table of Content

1. Introduction	4

1.1. Purpose	4

1.2. Scope	4

1.3. Terms, Definitions and Abbreviations	4

1.3.1. Definition of translation and rotation axes	5

2. Functional, Performance and Design requirements	5

2.1. General Requirements on Positioning systems	5

2.2. XYZ Sample positioner	6

2.3. Camera positioner	6

2.4. 5D positioners	7

2.5. 6D OAP positioner	8

2.6. 4D positioner	9

2.7. Linear positioners	10

2.8. Rotation positioners	10

2.9. 4D “dark” positioner	11

2.10. Motion controllers	12

2.10.1. General Requirements	12

2.10.2. Motion controller(s) MC1	12

2.10.3. Motion controller(s) MC2	12

2.10.4. Motion controller(s) MC3	13

3. Operational conditions	13

4. Delivery Requirements	13

5. Safety Requirements	14

6. Quality Requirements	14

6.1. General Quality Requirements	14

6.2. Specific Quality Requirements	14

# Introduction

## Purpose

This Requirements Specification Document (RSD) lists the technical requirements and constraints on positioning systems for RA4 experiments in experimental hall E1. The positioning systems will be used for various purposes, e.g. to hold and position a solid sample, align a focusing mirror, hold a camera, microscope, etc.

## Scope

This RSD contains all of the technical requirements: functional, performance, operational, transportation, safety and quality requirements for the following product: Positioning systems for E1 Endstations 2 (PBS number: E.E1.HXRS.POSS).

This product is a product Category B.

Category B is an Off-the-shelf Product with customization (e.g., product performance) that does not require any design modifications of the product. If there is no need in conducting specific tests regarding the product’s performances the Supplier shall verify all requirements defined by appropriate RSD within own outgoing inspection and testing procedure. All non-conformances (if any) must be addressed by the Supplier in a timely manner.

## Terms, Definitions and Abbreviations

For the purpose of this document, the following abbreviated terms are applied:

| Abbreviation | Meaning |
| --- | --- |
| ASCII | American Standard Code for Information Interchange |
| CA | Contracting Authority |
| ELI | Extreme Light Infrastructure |
| RA4 | Research Activity 4 (Applications in molecular, biomedical and materials sciences) |
| E1 | Experimental hall 1 |
| OAP | Off-axis parabola |
| PBS | Product Breakdown Structure |
| RSD | Requirements specification document |
| W x L x H | Width x Length x Height |

### Definition of translation and rotation axes

Figure 1: The schematic drawing for definition of translation and rotation axes.

# Functional, Performance and Design requirements

## General Requirements on Positioning systems

REQ-020688/A

The delivery of Positioning systems shall consist of:

- XYZ Sample positioner

- Camera positioner

- two 5D positioners

- 6D OAP positioner

- 4D positioner

- three Linear positioners

- two Rotation positioners:

  - RS1 for high vacuum

  - RS2 for ultra-high vacuum

- 4D “dark” positioner

- Motion controllers for positioners in 3. – 8.

NOTE:	All positioning systems shall be vacuum-compatible. Required level of vacuum compatibility for each system is defined individually in sections bellow.

REQ-020690/A

Cables (with connectors) between the positioners and the vacuum feedthroughs shall have length at least 600 mm.

NOTE:	Preferred connectors are either Sub-D 15 pin or connectors for LEMO feedthroughs. Feedthroughs shall NOT be included in the delivery.

REQ-020691/A

Cables shall have the same vacuum compatibility as positioners.

REQ-020692/A

All positioners shall have metric holes to mount a sample/device on the positioner.

## XYZ Sample positioner

REQ-020693/A

The XYZ Sample positioner shall be a stepper motor based positioner with 3 linear mutually orthogonal axes (x and y – horizontal, z – vertical).

REQ-020694/A

The XYZ Sample positioner shall have integrated limit switches to detect minimum and maximum positions of travel on each axis.

REQ-020695/A

Brackets and/or adapter plates to assemble the XYZ Sample positioner and mount it on a flat surface shall be included in the delivery.

REQ-020696/A

The XYZ Sample positioner shall have parameters listed in table n°1.

| Travel in x direction | 50 ± 1 mm |
| --- | --- |
| Travel in y direction | 100 ± 1 mm |
| Travel in z direction | 50 ± 1 mm |
| Load capacity | ≥ 1 kg |
| Resolution in full step (x and y direction) | ≤ 2.5 µm |
| Resolution in full step (z direction) | ≤ 1.25 µm |
| Total dimensions (xyz), including motors | less than 235 mm x 325 mm x 245 mm |
| Vacuum compatibility | 10-6 mbar or better |

Table n°1: Required parameters of the XYZ Sample positioner.

## Camera positioner

REQ-020697/A

The Camera positioner shall be a stepper motor based linear translation positioner.

REQ-020698/A

The Camera positioner shall have integrated limit switches to detect minimum and maximum positions of travel.

REQ-020699/A

It shall be possible to mount the Camera positioner on a standard optical breadboard.

NOTE:	An adapter plate can be used for mounting.

REQ-020700/A

The Camera positioner shall have parameters listed in table n°2.

| Travel range | 150 ± 2 mm |
| --- | --- |
| Load capacity | ≥ 7 kg |
| Resolution in full step | ≤ 2.5 µm |
| Total dimensions (xyz), including motors | Less than 325 mm x 80 mm x 45 mm |
| Vacuum compatibility | 10-6 mbar or better |

Table n°2: Required parameters of the Camera positioner.

## 5D positioners

REQ-020701/A

Two identical 5D positioners shall be delivered.

REQ-020702/A

Each 5D positioner shall be a piezo based positioner.

REQ-020703/A

Each 5D positioner shall have 3 linear and 2 angular motorized axes:

- x and y – horizontal axes

- z – vertical axis

- x – rotation around x axis

- y – rotation around y axis

REQ-020704/A

Each 5D positioner shall have parameters listed in table n°3

| Travel in x direction | ≥ 3mm |
| --- | --- |
| Travel in y direction | ≥ 3mm |
| Travel in z direction | ≥ 3mm |
| Angular travel x | ≥ 8 deg |
| Angular travel y | ≥ 8 deg |
| Load capacity | ≥ 1.2 kg |
| Minimum incremental motion (for linear axes) | ≤ 30 nm |
| Angular resolution | ≤ 0.7 rad |
| Total dimensions (xyz) | less than 120 mm x 85 mm x 42 mm |
| Vacuum compatibility | 10-8 mbar or better |

Table n°3: Required parameters of the 5D positioners.

## 6D OAP positioner

REQ-020705/A

The 6D OAP positioner shall be a piezo based positioner.

REQ-020706/A

The 6D OAP positioner shall consist of:

- 3 linear positioners: x and y – horizontal axes, z – vertical axis

- 2 goniometers: x and z – rotations around x and z axes

- 1 rotation positioner (360 deg): y – rotation around y axis

REQ-020707/A

The two goniometers shall have a common center of rotation.

REQ-020708/A

The 6D OAP positioner shall be equipped with sensors on all 6 axes to allow closed-loop operation.

REQ-020709/A

Brackets and/or adapter plates to assemble the 6D OAP positioner and mount it on a flat surface shall be included in the delivery.

REQ-020710/A

Proposed layout of the 6D OAP positioner shall be approved by the CA (see REQ-020753/A).

REQ-020711/A

System assembly with cabling and final testing shall be performed by the Supplier (within outgoing check of the Product, see REQ-020751/A).

REQ-020712/A

The 6D OAP positioner shall have parameters listed in table n°4

| Travel in x direction | ≥ 39 mm |
| --- | --- |
| Travel in y direction | ≥ 44 mm |
| Travel in z direction | ≥ 29 mm |
| Angular travel x | At least ± 5 deg |
| Angular travel y | 360 deg |
| Angular travel z | At least ± 5 deg |
| Load capacity | ≥ 150 g |
| Sensor resolution (for linear axes) | ≤ 2 nm |
| Repeatability measured over the complete travel range (for linear axes) | Better than ±70 nm |
| Angular sensor resolution (x and z) | ≤ 0.05 rad |
| Angular sensor resolution (y) | ≤ 0.3 rad |
| Total dimensions (xyz) | less than 98 mm x 120 mm x 115 mm |
| Vacuum compatibility | 10-6 mbar or better |

Table n°4: Required parameters of the 6D OAP positioner.

## 4D positioner

REQ-020713/A

The 4D positioner shall be a piezo based positioner.

REQ-020714/A

The 4D positioner shall consist of:

- 3 linear positioners: x and y – horizontal axes, z – vertical axis

- 1 rotation positioner (360 deg): z – rotation around z axis

REQ-020715/A

The 4D positioner shall be equipped with sensors on all 4 axes to allow closed-loop operation.

REQ-020716/A

Brackets and/or adapter plates to assemble the 4D positioner and mount it on a flat surface shall be included in the delivery.

REQ-020717/A

Proposed layout of the 4D positioner shall be approved by the CA (see REQ-020753/A).

REQ-020718/A

System assembly with cabling and final testing shall be performed by the Supplier (within outgoing check of the Product, see REQ-020751/A).

REQ-020719/A

The 4D positioner shall have parameters listed in table n°5

| Travel in x direction | ≥ 100 mm |
| --- | --- |
| Travel in y direction | ≥ 100 mm |
| Travel in z direction | ≥ 40 mm |
| Angular travel z | 360 deg |
| Load capacity | ≥ 300 g |
| Sensor resolution (for linear axes) | ≤ 2 nm |
| Repeatability measured over the complete travel range (for x and y axes) | Better than ± 155 nm |
| Repeatability measured over the complete travel range (for z axis) | Better than ± 80 nm |
| Angular sensor resolution (for rotation positioner) | ≤ 0.3 rad |
| Total dimensions (xyz) | less than 155 mm x 155 mm x 140 mm |
| Vacuum compatibility | 10-6 mbar or better |

Table n°5: Required parameters of the 4D sample positioner.

## Linear positioners

REQ-020720/A

Three identical Linear positioners shall be delivered.

REQ-020721/A

Each Linear positioner shall be a piezo based linear translation positioner.

REQ-020722/A

Each Linear positioner shall be equipped with a sensor to allow closed-loop operation.

REQ-020723/A

Each Linear positioner shall have parameters listed in table n°6.

| Travel range | 49 ± 1 mm |
| --- | --- |
| Load capacity | ≥ 3 kg |
| Sensor resolution | ≤ 2 nm |
| Repeatability measured over the complete travel range | Better than ± 80 nm |
| Total dimensions (xyz) | Less than 76 mm x 25 mm x 11 mm |
| Vacuum compatibility | 10-6 mbar or better |

Table n°6: Required parameters of the Linear positioners.

## Rotation positioners

REQ-020724/A

Two rotation positioners – RS1 and RS2 – shall be delivered.

REQ-020725/A

The RS1 positioner shall be compatible with vacuum level 10-6 mbar or better.

REQ-020726/A

The RS2 positioner shall be compatible with vacuum level 10-8 mbar or better.

REQ-020727/A

Each Rotation positioner shall be piezo based.

REQ-020728/A

Each Rotation positioner shall be equipped with a sensor to allow closed-loop operation.

REQ-020729/A

Each Rotation positioner shall have parameters listed in table n° 7.

| Angular travel range | 360 deg |
| --- | --- |
| Load capacity | ≥ 2.5 kg |
| Sensor resolution | ≤ 0.3 rad |
| Total dimensions (xyz) | Less than 58 mm x 58 mm x 14 mm |

Table n°7: Required parameters of the Rotation positioners.

## 4D “dark” positioner

REQ-020730/A

The 4D “dark” positioner shall be a piezo based positioner.

REQ-020731/A

The 4D “dark” positioner shall consist of:

- 3 linear positioners: x and y – horizontal axes, z – vertical axis

- 1 rotation positioner (360 deg): z – rotation around z axis

REQ-020732/A

The 4D “dark” positioner shall be equipped with sensors on all 4 axes to allow closed-loop operation.

REQ-020733/A

The sensors on the 4D “dark” positioner shall not use light with wavelength below 1.2 m to detect the position.

NOTE:	Other types of sensors, e.g. resistive or inductive, can be used.

REQ-020734/A

Brackets and/or adapter plates to assemble the 4D “dark” positioner and mount it on a flat surface shall be included in the delivery.

REQ-020735/A

The 4D “dark” positioner shall have parameters listed in table n°8.

| Travel in x direction | ≥ 20 mm |
| --- | --- |
| Travel in y direction | ≥ 20 mm |
| Travel in z direction | ≥ 12 mm |
| Angular travel z | 360 deg |
| Encoded angular travel | ≥ 310 deg |
| Load capacity | ≥ 140 g |
| Sensor resolution (for linear axes) | ≤ 250 nm |
| Repeatability (for linear axes) | Better than 3 m |
| Absolute accuracy (for linear axes) | Better than 1.5% of travel range |
| Angular sensor resolution (for rotation positioner) | ≤ 0.11 mrad |
| Absolute accuracy (for rotation) | Better than 2% of travel range |
| Total dimensions (xyz) | less than 48 mm x 48 mm x 74 mm |
| Vacuum compatibility | 10-8 mbar or better |

Table n°8: Required parameters of the 4D “dark” positioner.

## Motion controllers

### General Requirements

REQ-020736/A

The delivery of motion controllers shall consist of:

- MC1 – motion controller (or controllers) for 5D positioners

- MC2 – motion controller (or controllers) for 6D OAP positioner, 4D positioner, Linear positioners and Rotation positioners

- MC3 – motion controller (or controllers) for 4D “dark” positioner

NOTE:	19” rack-mounted controllers are preferred.

REQ-020737/A

The motion controllers shall be delivered with EU compatible power supplies.

REQ-020738/A

The motion controllers shall be delivered with cables between vacuum feedthroughs and the controller.

REQ-020739/A

The motion controllers shall have RJ45 or RS232 socket that allows communication and control from a remote host computer. This control shall be possible through one of the following protocols:

- ASCII based command protocol (protocol is documented);

- Linux 64bit compatible C++ libraries (libraries are part of the delivery).

NOTE:	The RJ45 socket can be provided by extender/convertor or by mounting on a DIN rail.

### Motion controller(s) MC1

REQ-020740/A

The MC1 shall have at least 10 channels to control two 5D positioners.

REQ-020741/A

The MC1 shall operate in open-loop.

### Motion controller(s) MC2

REQ-020742/A

The MC2 shall be compatible with these positioners:

- 6D OAP positioner

- 4D positioner

- Linear positioners

- Rotation positioners

REQ-020743/A

The MC2 shall have 15 channels to control 15 motion axes.

REQ-020744/A

The MC2 shall use feedback from the positioner’s sensor to control the position (closed-loop operation).

### Motion controller(s) MC3

REQ-020745/A

The MC3 shall have at least 4 channels to control the 4D “dark” positioner.

REQ-020746/A

The MC3 shall use feedback from the positioner’s sensor to control the position (closed-loop operation).

# Operational conditions

REQ-020757/A

The Supplier and the CA shall agree on the cleaning method to clean devices that are not placed in vacuum (e.g. controllers) without decreasing the devices’ performance and to avoid contamination when used in a Class 7 clean environment according to ISO 14644.

REQ-020758/A

All elements of the positioning systems shall be suitable for operation at temperature 18 – 25 °C.

# Delivery Requirements

REQ-020747/A

The transportation to the final destination shall be conducted by the Supplier.

REQ-020748/A

The technologies and instruments shall be delivered in protective package preventing damage and contamination and a minimum of two plies separate clean packaging.

# Safety Requirements

REQ-020749/A

The Supplier shall supply a Declaration of Conformity or any other equivalent document legally recognized and accepted in the Czech Republic for each product type if the appropriate legislation determines the Supplier's obligation to have a Declaration of Conformity (or the equivalent document) for the purposes of a Device sale in the Czech Republic to fulfil the requirements of 2001/95/EC directive or applicable Czech law.

# Quality Requirements

## General Quality Requirements

REQ-020750/A

The Supplier shall provide the Product User Manual as part of the delivered Device (positioning systems). The Manual shall include the instructions and descriptions regarding the following procedures:

- transport, handling, storage, installation and cleaning;

- safe operation and maintenance procedures.

REQ-020751/A

The Supplier shall provide information on execution of outgoing check of the Product. At least this information shall comprise declaration about execution of outgoing check and declaration of conformity with technical requirements defined by the product RSD and completeness of the Product.

NOTE: Alternatively the Supplier might provide the CA with information (e.g.: catalogue/technical data sheets, product manuals or other similar documentation) subject to the condition that such documentation is detailed enough to prove meeting all requirements stipulated herein.

REQ-020752/A

The Supplier shall establish and maintain a non-conformance control system compatible with ČSN EN ISO 9001.

## Specific Quality Requirements

REQ-020753/A

The Supplier shall provide a final suggestion of each positioning system configuration that is to be accepted by the CA before manufacturing starts.

REQ-020754/A

The unit system of the mounting holes, sizes and layout shall be SI base units (metric) or other units derived from them.