|  |  |  |  |
| --- | --- | --- | --- |
| **Confidentiality Level** | *BL - Restricted for internal use* | **TC ID / Revision** | 00155562/C |
| **Document Status** | *Document Released* | **Document No.** | N/A |
| **WBS code** | *5.4 - RP5 Laser plasma and high-energy-density physics* |
| **PBS code** | *E.E3.P3DG.3* |
| **Project branch** | *Engineering & Scientific documents (E&S)* |
| **Document Type** | *Specification (SP)* |
| ***[RSD product category A]******Dpt. 92******Scientific grade cameras and spectrometers*** ***TP17\_063***logoKeywordsImaging spectrometers, scientific cameras, technical requirements, parameters  |
|  | **Position** | **Name** |
| **Responsible person** | Junior Researcher RP5 | Singh Sushil Kumar |
| **Prepared by** | Junior Researcher RP5Junior Researcher RP5 | Singh Sushil KumarBohlin Hannes |

|  |  |  |  |
| --- | --- | --- | --- |
| RSS TC ID/revision | RSS - Date of Creation | RSS - Date of Last Modification | Systems Engineer |
| 012755/A.001 | 03.08.2017 19:41 | 03.08.2017 19:43 | Aleksei Kuzmenko |
| 012755/A.002 | 10.08.2017 19:59 | 10.08.2017 20:00 | Aleksei Kuzmenko |
| 012755/A.003 | 17.08.2017 12:19 | 17.08.2017 12:22 | Aleksei Kuzmenko |

|  |
| --- |
| ****Reviewed By**** |
| Name (Reviewer) | Position | Date | Signature |
| Batheja Deepak Kumar | Senior Researcher RP5 | NOTICE (RSD product category A) |
| Ladislav Půst | Manager installation of technology | NOTICE (RSD product category A) |
| Pavel Korouš | Chief Engineer | NOTICE (RSD product category A) |
| Petr Pivoňka | Control System Engineer | NOTICE (RSD product category A) |
| Roman Kuřátko | Facility Manager | NOTICE (RSD product category A) |
| Veronika Olšovcová | Safety Coordinator | NOTICE (RSD product category A) |
| Viktor Fedosov | SE & Planning group leader;Quality Manager | NOTICE (RSD product category A) |
| Weber Stefan Andreas | RP5 / RP6 Team Leader | NOTICE (RSD product category A) |

|  |
| --- |
| ****Approved by**** |
| Name (Approver) | Position | Date | Signature |
| Georg Korn | Science and Technology Manager, Scientific coordinator of RP2-6 |  |  |

|  |
| --- |
| ****Revision History / Change Log**** |
| Change No. | Made by | Date | Change description, Pages, Chapters | TC rev. |
| 1 | Singh Sushil Kumar,Bohlin Hannes | 03.08.2017 | RSD draft creation | A |
| 2 | A.Kuzmenko,Singh Sushil Kumar | 10.08.2017 | RSD update, version for internal review | B |
| 3 | A.Kuzmenko | 16.08.2017 | RSD update, final version for approval | C |

Table of Content

[1. Introduction 4](#_Toc490662520)

[1.1. Purpose 4](#_Toc490662521)

[1.2. Scope 4](#_Toc490662522)

[1.3. Terms, Definitions and Abbreviations 4](#_Toc490662523)

[1.4. References to standards 5](#_Toc490662524)

[2. Functional, Performance and Design requirements 6](#_Toc490662525)

[2.1. Tender part A - Imaging spectrometers 6](#_Toc490662526)

[2.1.1. General requirements for all the spectrometers 6](#_Toc490662527)

[2.1.1.1. Software and computer interface requirements 6](#_Toc490662528)

[2.1.1.2. Power supply 6](#_Toc490662529)

[2.1.1.3. Turret interface 6](#_Toc490662530)

[2.1.2. Turret and grating requirements for all the spectrometers 7](#_Toc490662531)

[2.1.2.1. General requirements 7](#_Toc490662532)

[2.1.2.2. Turret #1 7](#_Toc490662533)

[2.1.2.3. Turret #2 7](#_Toc490662534)

[2.1.3. Imaging Spectrometer #1 8](#_Toc490662535)

[2.1.3.1. Focal length 8](#_Toc490662536)

[2.1.3.2. Wavelength accuracy 8](#_Toc490662537)

[2.1.3.3. Wavelength repeatability 8](#_Toc490662538)

[2.1.3.4. Input port 8](#_Toc490662539)

[2.1.3.5. Output ports 8](#_Toc490662540)

[2.1.3.6. Optics 8](#_Toc490662541)

[2.1.4. Imaging Spectrometer #2 9](#_Toc490662542)

[2.1.4.1. Focal length 9](#_Toc490662543)

[2.1.4.2. Wavelength accuracy 9](#_Toc490662544)

[2.1.4.3. Wavelength repeatability 9](#_Toc490662545)

[2.1.4.4. Input port 9](#_Toc490662546)

[2.1.4.5. Output ports 9](#_Toc490662547)

[2.1.4.6. Optics 9](#_Toc490662548)

[2.1.5. Imaging Spectrometer #3 10](#_Toc490662549)

[2.1.5.1. Focal length 10](#_Toc490662550)

[2.1.5.2. Wavelength accuracy 10](#_Toc490662551)

[2.1.5.3. Wavelength repeatability 10](#_Toc490662552)

[2.1.5.4. Input port 10](#_Toc490662553)

[2.1.5.5. Output ports 10](#_Toc490662554)

[2.1.5.6. Optics 10](#_Toc490662555)

[2.1.6. Camera for the Spectrometers 11](#_Toc490662556)

[2.1.6.1. General requirements 11](#_Toc490662557)

[2.2. Tender part B – sCMOS scientific imaging camera 12](#_Toc490662558)

[2.2.1. General requirements 12](#_Toc490662559)

[2.2.1.1. Software and computer interface requirements 12](#_Toc490662560)

[2.3. Tender part C – Large field-of-view imaging camera 13](#_Toc490662561)

[2.3.1. General requirements 13](#_Toc490662562)

[2.3.1.1. Software and computer interface requirements 13](#_Toc490662563)

[3. Delivery requirements 14](#_Toc490662564)

[4. Safety Requirements 14](#_Toc490662565)

[5. Quality Requirements 15](#_Toc490662566)

[5.1. General Quality Requirements 15](#_Toc490662567)

[5.2. Specific Quality requirements 15](#_Toc490662568)

# Introduction

## Purpose

This Requirements Specification Document (RSD) lists the technical requirements and constraints on products being purchased for the department 92 of ELI Beamlines project.

## Scope

The RSD contains all of the technical requirements: functional, performance and design, delivery, safety and quality requirements for the following products: **Imaging spectrometers and scientific grade cameras** (*PBS code: E.E3.P3DG.3*).

These products are in product **Category A** according to the ELI Beamlines RSD categories of products. The Category A is an Off-the-shelf Product without necessity of modifications and necessity to be subjected to a verification program (review of design, inspection and testing) for ELI applications by the actual project specifications. All verification activities performing by a supplier shall be executed in accordance with the supplier’s plan of outgoing inspection and tests. Internal Acceptance Procedure of the product Category A shall be established and applied before the product implementation (operation phase).

## Terms, Definitions and Abbreviations

For the purpose of this document, the following abbreviated terms are applied:

| **Abbreviation** | **Meaning** |
| --- | --- |
| AC | Alternating current |
| ADC | Analog-to-Digital Converter |
| CA | Contracting Authority (Institute of Physics AV CR, v. v. i.) |
| CCD | Charge-Coupled Device |
| CMOS | Complementary Metal Oxide Semiconductor |
| ELI | Extreme Light Infrastructure |
| EM | Electron multiplication |
| EMCCD | Electron Multiplying CCD |
| GigE | Gigabit Ethernet |
| PBS | Product Breakdown Structure |
| RSD | Requirements Specification Document |
| sCMOS | Scientific CMOS technology |
| SDK | Software Development Kit |
| TTL | Transistor–Transistor Logic |
| USB | Universal Serial Bus |
| UV | UltraViolet |

For the purpose of this document, the following terms and definitions are applied:

|  |  |
| --- | --- |
| **Term** | **Definition** |
| Wavelength accuracy of a spectrometer | Wavelength accuracy of a spectrometer is defined as the difference between the predicted position of an emission line (as predicted by the spectrometer control software) and the measured peak position of the atomic emission line. |
| Wavelength reproducibility of a spectrometer | Wavelength reproducibility is defined as the maximum deviation in measurement of centroid of a particular spectral line during multiple scans (20 times). The wavelength range of the spectrometer shall be set moving from low wavelength to high wavelength (10 times) and moving from high wavelength to low wavelength (10 times).  |

## References to standards

If this document includes references to standards or technical documents the CA allows/permits also another equal solution to be offered. If the Supplier offers another equal solution the CA shall not reject its bid, once the Supplier by appropriate means in the bid proves that the offered supplies, services or works meet in an equivalent manner the requirements including references to standards or technical documents.

# Functional, Performance and Design requirements

## Tender part A - Imaging spectrometers

### General requirements for all the spectrometers

#### Software and computer interface requirements

REQ-021493/A

The Supplier shall provide the required software for the spectrometers control, data acquisition and display.

REQ-021494/A

The software shall be able to run on Windows 10 operating system (or newer versions).

REQ-021598/A

The Supplier shall provide a software development kit (SDK) allowing programmatic control of full range of spectrometers settings and regimes and undepreciated raw image data acquisition.

*NOTE 1: The SDK shall be compatible with Windows 10 and Centos 7 operating systems (or newer versions).*

*NOTE 2: The* *SDK shall be delivered in form of 64-bit linkable library binaries (.dll for Windows, .so for Linux), C/C++ header includes and full-scale documentation of all public library objects.*

REQ-021495/A

The CA shall be able to install the software in any number of computers without incurring any additional costs.

REQ-021496/A

The control interface of the spectrometers shall be compatible with USB 3.0.

REQ-021498/A

The Supplier shall provide a cable in the range of 2 m to 3 m long for the control of the spectrometer.

#### Power supply

REQ-021497/A

All the spectrometers shall have power supply compatible with input Voltage in the range of 210–245 V AC at 50–60 Hz.

#### Turret interface

REQ-021508/A

The turret control of all the spectrometers shall be motorized.

REQ-021509/A

All the spectrometers and the turrets shall be designed such that the CA can replace/install a different turret in the future.

### Turret and grating requirements for all the spectrometers

#### General requirements

REQ-021499/A

The Supplier shall provide 2 turrets that shall comply with the following specifications:

* **Turret #1** shall be compatible with imaging spectrometer #1;
* **Turret #2** shall be compatible with imaging spectrometers #2 and #3 (see chapters 2.1.4 and 2.1.5).

REQ-021500/A

The gratings shall be plane ruled reflectance grating.

#### Turret #1

REQ-021501/A

The turret shall house two gratings.

REQ-021502/A

The gratings shall have the following specifications:

1. **Grating 1A** shall have a groove density of 150 lines/mm, blazed for 500 nm;
2. **Grating 1B** shall have a groove density of 300 lines/mm, blazed for 422 nm.

REQ-021503/A

*The gratings* ***1A*** *and* ***1B*** *shall be coated with Aluminum coating.*

#### Turret #2

REQ-021504/A

The turret shall house three gratings.

REQ-021505/A

The gratings shall have the following specifications:

1. **Grating 2A** shall have a groove density of 150 lines/mm, blazed for 500 nm;
2. **Grating 2B** shall have a groove density of 300 lines/mm, blazed for 750 nm;
3. **Grating 2C** shall have a groove density of 1200 lines/mm, blazed for 1000 nm.

REQ-021600/A

*The gratings* ***2A, 2B*** *and* ***2C*** *shall be coated with Aluminum coating.*

### Imaging Spectrometer #1

#### Focal length

REQ-021506/A

The spectrometer shall be designed in Czerny Turner configuration.

The focal length of both the curved mirrors shall be between 175 ± 25 mm.

REQ-021507/A

The focusing optics and grating shall be such that the f-number at the input shall be between f/3.5 and f/4.1.

#### Wavelength accuracy

REQ-021510/A

The wavelength accuracy of the spectrometer #1 shall be equal or less than 0.25 nm, when used with a grating with groove density 1200 l/mm.

#### Wavelength repeatability

REQ-021511/A

The wavelength repeatability of the spectrometer #1 shall be less than 75 pm, when used with a grating with groove density 1200 l/mm.

#### Input port

REQ-021512/A

The spectrometer shall have one input port with motorized slit assembly.

#### Output ports

REQ-021513/A

The spectrometer shall have single output port.

REQ-021514/A

The output port shall be able to provide an imaging aperture of 25 x 10 mm.

REQ-021515/A

The Supplier shall provide a mechanical slit that can be attached to the output port.

#### Optics

REQ-021516/A

The optics within the spectrometer shall be Al+MgF2 coated to enhance transmission in ultra violet region.

### Imaging Spectrometer #2

#### Focal length

REQ-021517/A

The spectrometer shall be designed in Czerny Turner configuration.

The focal length of both the curved mirrors shall be between 500 ± 50 mm.

REQ-021518/A

The focusing optics and grating shall be such that the f-number at the input shall be between f/6 and f/7.

#### Wavelength accuracy

REQ-021521/A

The wavelength accuracy of the spectrometer #2 shall be equal or less than 0.2 nm, when used with a grating with groove density 1200 l/mm.

#### Wavelength repeatability

REQ-021522/A

The wavelength repeatability of the spectrometer #2 shall be less than 40 pm, when used with a grating with groove density
1200 l/mm.

#### Input port

REQ-021523/A

The spectrometer #2 shall have one input port with motorized slit assembly.

#### Output ports

REQ-021524/A

The spectrometer #2 shall have 2 output ports. The output port shall be selectable with a motorized mirror inside the spectrometer.

REQ-021525/A

Each of the spectrometer #2 output port shall be able to provide an imaging aperture of 30 x 14 mm.

REQ-021526/A

The Supplier shall provide a mechanical slit that can be attached to one of the output ports.

#### Optics

REQ-021527/A

The optics within the spectrometer #2 shall be silver coated to enhance transmission in Visible and Near Infra-red region.

### Imaging Spectrometer #3

#### Focal length

REQ-021528/A

The spectrometer shall be designed in Czerny Turner configuration.

The focal length of both the curved mirrors shall be between 750 ± 50 mm.

REQ-021529/A

The focusing optics and grating shall be such that the f-number at the input shall be between f/9.0 and f/10.5.

#### Wavelength accuracy

REQ-021532/A

The wavelength accuracy of the spectrometer #3 shall be equal or less than 0.2 nm, when used with a grating with groove density 1200 l/mm.

#### Wavelength repeatability

REQ-021533/A

The wavelength repeatability of the spectrometer #3 shall be equal or less than 10 pm, when used with a grating with groove density 1200 l/mm.

#### Input port

REQ-021534/A

The spectrometer #3 shall have one input port with motorized slit assembly.

#### Output ports

REQ-021535/A

The spectrometer #3 shall have 2 output ports. The output port shall be selectable with a motorized mirror inside the spectrometer.

REQ-021536/A

Each of the output port shall be able to provide an imaging aperture of 30 x 14 mm.

REQ-021537/A

The Supplier shall provide a mechanical slit that can be attached to one of the output ports.

#### Optics

REQ-021538/A

The optics within the spectrometer shall be silver coated to enhance transmission in Visible and Near Infra-red region.

### Camera for the Spectrometers

#### General requirements

REQ-021539/A

The parameters of all the cameras for the spectrometers (see chapters 2.1.3, 2.1.4 and 2.1.5) shall correspond to the requirements given in table 1 below.

| **Item N°** | **Parameters** | **Minimum Requirements** |
| --- | --- | --- |
| 1.1 | Anti-reflection coating | UV coating (QE ≥ 35 % at 250 nm) |
| 1.2 | Sensor type | Back illuminated EMCCD with fringe suppression technology |
| 1.3 | Fringe suppression | Intensity fluctuation resulting from interference should be reduced to max. ±20% (peak-to-peak) in the wavelength range 250 nm – 1000 nm |
| 1.4 | Pixel format | 1600 x 400 pixels |
| 1.5 | Pixel size | 16 x 16 μm |
| 1.6 | Sensor area | 25.6 x 6.4 mm |
| 1.7 | Linearity | ≥ 99 % at 1MHz readout rate |
| 1.8 | Quantum efficiency | ≥ 95 % at maximum (peak located in 500-800 nm range), ≥ 35 % at 250 nm |
| 1.9 | Spectrometric well capacity | ≥ 200,000 electrons |
| 1.10 | Readout noise | ≤ 25 electrons/pixel/s (typical) in EM mode at 1MHz readout rate |
| 1.11 | Dark current | < 0.01 electrons/pixel/s |
| 1.12 | Clock-induced charge | < 0.02 electrons/pixel/frame |
| 1.13 | Binning | Should be selectable |
| 1.14 | Readout interface | USB version 2.0 or higher, or GigE |
| 1.15 | Trigger input | Fiber optic or TTL |
| 1.16 | ADC resolution | ≥ 16 bit |
| 1.17 | Cooling | Thermoelectric |

**Table 1**: Functional and performance parameters of the cameras for spectrometers.

##### Camera interface requirements

REQ-021614/A

The Supplier shall provide the relevant hardware required to install the camera on any of the spectrometers described in the previous sections.

## Tender part B – sCMOS scientific imaging camera

### General requirements

REQ-021540/A

The parameters of all the sCMOS scientific imaging cameras shall correspond to the requirements given in table 2 below.

| **Item N°** | **Parameters** | **Minimum Requirements** |
| --- | --- | --- |
| 2.1 | Sensor type | sCMOS |
| 2.2 | Pixel format | 2048 × 2048 pixels |
| 2.3 | Pixel size | 6.5 x 6.5 µm |
| 2.4 | Sensor area | 13.3 × 13.3 mm |
| 2.5 | Linearity | ≥ 99 % |
| 2.6 | Quantum efficiency | ≥ 80 % (peak) |
| 2.7 | Full well capacity | ≥ 30,000 electrons |
| 2.8 | Readout noise | < 1.0 electrons (median) |
| 2.9 | Dark current | < 0.02 electrons/pixel/s |
| 2.10 | Dynamic range | ≥ 33,000:1 |
| 2.11 | Readout interface | USB3.0 |
| 2.12 | ADC resolution | ≥ 16 bit |
| 2.13 | Trigger input | Fiber optic or TTL |
| 2.14 | Optical interface | C-mount |
| 2.15 | Cooling | Thermoelectric |

**Table 2**: Functional and performance parameters of the sCMOS scientific imaging camera.

#### Software and computer interface requirements

REQ-021615/A

The Supplier shall provide the required software for camera control, data acquisition and display.

REQ-021616/A

The software shall be able to run on Windows 10 operating system (or newer versions).

REQ-021617/A

The Supplier shall provide a software development kit (SDK) allowing programmatic control of full range of camera settings and regimes and undepreciated raw image data acquisition.

*NOTE 1: The SDK shall be compatible with Windows 10 and Centos 7* *operating systems (or newer versions).*

*NOTE 2: The SDK shall be delivered in form of 64-bit linkable library binaries (.dll for Windows, .so for Linux), C/C++ header includes and full-scale documentation of all public library objects.*

REQ-021618/A

The CA shall be able to install the software in any number of computers without incurring any additional costs.

REQ-021619/A

The Supplier shall provide a cable between 2 m to 3 m long for the control of the camera.

## Tender part C – Large field-of-view imaging camera

### General requirements

REQ-021541/A

The parameters of all the large field-of-view imaging cameras shall correspond to the requirements given in table 3 below.

| **Item N°** | **Parameters** | **Minimum Requirements** |
| --- | --- | --- |
| 3.1 | Sensor type | CCD or sCMOS |
| 3.2 | Pixel format | ≥ 4096 x 4096 pixels |
| 3.3 | Pixel size | 9 x 9 µm |
| 3.4 | Sensor area | ≥ 36.8 x 36.8 mm |
| 3.5 | Linearity | ≥ 99 % |
| 3.6 | Quantum efficiency | ≥ 60 % (peak) |
| 3.7 | Full well capacity | ≥ 90,000 |
| 3.8 | Readout noise | < 10 electrons |
| 3.9 | Dark current | < 0.02 electrons/pixel/s |
| 3.10 | Dynamic range | ≥ 9,800:1 |
| 3.11 | Readout interface | USB or Ethernet |
| 3.12 | ADC resolution | ≥ 16 bit |
| 3.13 | Cooling | Thermoelectric |

**Table 3**: Functional and performance parameters of the large field-of-view imaging camera.

#### Software and computer interface requirements

REQ-021620/A

The Supplier shall provide the required software for camera control, data acquisition and display.

REQ-021621/A

The software shall be able to run on Windows 10 operating system (or newer versions).

REQ-021622/A

The Supplier shall provide a software development kit (SDK) allowing programmatic control of full range of camera settings and regimes and undepreciated raw image data acquisition.

*NOTE 1: The SDK shall be compatible with Windows 10 and Centos 7 operating systems (or newer versions).*

*NOTE 2: The SDK shall be delivered in form of 64-bit linkable library binaries (.dll for Windows, .so for Linux), C/C++ header includes and full-scale documentation of all public library objects.*

REQ-021623/A

The CA shall be able to install the software in any number of computers without incurring any additional costs.

REQ-021624/A

The Supplier shall provide a cable between 2 m to 3 m long for the control of the camera.

# Delivery requirements

REQ-021602/A

The transportation to the final destination shall be conducted by the Supplier.

# Safety Requirements

REQ-021603/A

The Supplier shall supply a Declaration of Conformity or any other equivalent document legally recognized and accepted in the Czech Republic for each product type if the appropriate legislation determines the Supplier's obligation to have a Declaration of Conformity (or the equivalent document) for the purposes of a Device sale in the Czech Republic to fulfil the requirements of 2001/95/EC directive or applicable Czech law.

# Quality Requirements

## General Quality Requirements

REQ-021604/A

The Supplier shall provide the Product User Manual as part of the delivered Product. The Manual shall include the instructions and descriptions regarding the following procedures:

* transport, handling and storage;
* installation, cleaning and calibration;
* user manual for the software;
* safe operation and maintenance procedures.

*NOTE 1: The Product technical data sheet shall be part of manual.*

*NOTE 2: The scope of the User Manual shall be agreed with the CA before Product delivery.*

REQ-021605/A

The Supplier shall establish and maintain a non-conformance control system compatible with ČSN EN ISO 9001 (equivalent to EN ISO 9001).

## Specific Quality requirements

REQ-021606/A

In case of the spectrometer repair by the Supplier within the validity of warranty, the Supplier shall recalibrate and verify the spectrometer once the repair is completed. The results of these processes shall be provided to the CA.

REQ-021607/A

All the spectrometers shall be delivered with documented results of the Supplier’s outgoing check (e.g. test protocols, inspection report and etc.).

*NOTE 1: These results shall demonstrate that the spectrometers comply with technical requirements* *stipulated herein.*

*NOTE 2: The test equipment used for performing the tests shall have valid metrological confirmation.*