| Confidentiality Level |  | BL - Restricted for internal use | TC ID / Revision | 00158464/C |  |
| --- | --- | --- | --- | --- | --- |
| Document Status |  | Document Released | Document No. | N/A |  |
| WBS code |  | 3.4 – L4 system |  |  |  |
| PBS code |  | RA1.L4.CMP1.CIS |  |  |  |
| Project branch |  | Engineering & Scientific documents (E&S) |  |  |  |
| Document Type |  | Specification (SP) |  |  |  |
| [RSD product category C] Annex No. 2 –Detailed technical specifications Integrated system of vacuum chambers, optomechanical mounts and controls of the L4 Compressor Image relay System (CIS) [TP17-044] Keywords N/A |  |  |  |  |  |
|  | Position |  |  |  | Name |
| Responsible person | Scientific coordinator of laser technology (RP1) |  |  |  | Bedřich Rus |
| Prepared by | Senior Technology Expert Scientific Coordinator of laser technology (RP1) Lead Laser Control and Timing Systems Specialist Senior Laser Optical Specialist |  |  |  | Mike Griffiths Bedřich Rus Jack Naylon Daniel Kramer |

| RSS TC ID/revision | RSS - Date of Creation | RSS - Date of Last Modification | Systems Engineer |
| --- | --- | --- | --- |
| 008547/A.001 | 08.11.2017 15:16 | 08.11.2017 15:18 | Aleksei Kuzmenko |
| 008547/A.002 | 10.11.2017 14:27 | 10.11.2017 14:29 | Aleksei Kuzmenko |
| 008547/A.003 | 27.11.2017 13:24 | 27.11.2017 13:24 | Aleksei Kuzmenko |

| Reviewed By |  |  |  |
| --- | --- | --- | --- |
| Name (Reviewer) | Position | Date | Signature |
| Alice Hamalová | Clean room specialist | N O T I C E |  |
| Jakub Janďourek | Infrastructure technology coordinator | N O T I C E |  |
| Ladislav Půst | Manager installation of technology | N O T I C E |  |
| Lukáš Brabec | Senior Vacuum Specialist |  |  |
| Martin Laub | Head of Mechanical Engineering |  |  |
| Pavel Bakule | Deputy RP1 Leader |  |  |
| Pavel Korouš | Chief Engineer |  |  |
| Pavel Trojek | Senior Laser Specialist |  |  |
| Roman Kuřátko | Facility Manager | N O T I C E |  |
| Veronika Olšovcová | Safety Coordinator |  |  |
| Viktor Fedosov | SE & Planning group leader; Quality Manager |  |  |

| Approved by |  |  |  |
| --- | --- | --- | --- |
| Name (Approver) | Position | Date | Signature |
| Bedřich Rus | Scientific Coordinator of Laser Technology (RP1) |  |  |

| Revision History / Change Log |  |  |  |  |
| --- | --- | --- | --- | --- |
| Change No. | Made by | Date | Change description, Pages, Chapters | TC rev. |
| 1 | B.Rus | 17.10.2017 | RSD draft creation | A |
| 2 | B.Rus A.Kuzmenko | 10.11.2017 | RSD update, version for internal review | B |
| 3 | B.Rus A.Kuzmenko | 27.11.2017 | RSD update, final version for approval | C |

# Content

1.	Introduction	4

1.1. References to standards	5

2.	Scope of work	6

3.	Arrangement of the CIS components and integration with L4 laser system	7

4.	Specific design of CIS1A unit	11

4.1. CIS1A Optical Configuration	11

4.2. CIS1A Vacuum Chamber	11

4.3. CIS1A Optical Support Chassis	12

4.4. CIS1A Optomechanical Mounts	14

4.5. CIS1A Laser Shutter	15

4.6. Interconnecting tube between CIS1A and CIS1B chambers	16

5.	Specific design of CIS1B unit	17

5.1. CIS1B Optical Configuration	17

5.2. CIS1B Vacuum Chamber	17

5.3. CIS1B Optical Support Chassis	18

5.4. CIS1B Optomechanical Mounts	19

6.	Specific design of CIS2 unit	21

6.1. CIS2 Optical Configuration	21

6.2. CIS2 Vacuum Chamber	21

6.3. CIS2 Optical Support Chassis	22

6.4. CIS2 Optomechanical Mounts	23

7.	General Requirements for Vacuum Chamber Design and Manufacture	24

8.	General Requirements for Vacuum Equipment Design and Supply	26

8.1. CIS vacuum circuit and its connection to ELI-Beamlines central vacuum	26

8.2. Cable trays	29

9.	CIS Control Systems	30

9.1. Vacuum Control System	30

9.2. Motion Control System	37

10.	Vacuum flange schedule	42

10.1. CIS1A chamber	42

10.2. CIS1B chamber	45

10.3. CIS2 chamber	47

11.	Summarised design and functional requirements for the CIS	49

11.1. Layout and optomechanical configuration	49

11.2. Optomechanical mounts	51

11.3. Vacuum components	52

11.4. Optical Support Chassis	55

11.5. Control systems	56

11.6. Packaging and Transportation requirements	57

11.7. Testing, inspection, documentation and warranty	57

11.8. Safety Requirements	58

12.	List of Abbreviations	60

13.	Reference drawings and diagrams	61

13.1. Drawing Package	61

# Introduction

This Requirements Specification Document (RSD) states and describes the technical requirements for the integrated assembly of vacuum chambers, optomechanical systems and electronic controls of the Compressor Image relay System (CIS). The CIS is an integral part of the overall L4 laser system of ELI-Beamlines and is registered in the PBS database under the following PBS code: RA1.L4.CMP1.CIS.

The CIS system connects the final amplifier of the L4 laser system, located in the L4b hall, with the 10 PW vacuum compressor located in the L4c hall. The CIS design as described in this document has been closely coordinated with the L4 laser and with the design of the 10 PW compressor (produced by FZU).

An overall view of the CIS components (comprising units CIS1A, CIS1B and CIS2) in relation to the L4 Final Amplifier, the L4 Compressor and the building structure is shown in Figure 1.

The Compressor Imaging System (CIS) provides the following functions:

- To image the output beam from the L4 laser and direct it along the required geometric path to reach the input of the compressor in L4c one floor below.

- To transport the beam entirely under vacuum.

- To pass full-aperture and sub-aperture samples of the uncompressed beam to diagnostic packages in L4b.

- To allow the L4 laser output beam to be switched on demand into a beam dump in L4b laser hall.

- To provide an optional path to direct the uncompressed beam into experimental hall L3.

Figure 1: Overall view of the L4 10 PW compressor and of the Compressor Image relay System (CIS). Also shown in the figure is the laser optical chamber VC2 which is not part of CIS.

This document is comprised of a descriptive text that explains the purpose, function and arrangement of the equipment to be supplied and is supported by further sections that provide a summary of requirements and technical details of material that comprises the scope of supply under the contract; all materials shall be compliant with the specified details.

This specification addresses the interfaces with the L4 laser chain, the L4 10 PW compressor, the ELI-Beamlines building and the L4 control systems. This RSD also acts as a parent document for any technical requirements specified in lower level design documentation.

The drawings attached to this RSD provide information on essential features of the CIS1A, CIS1B andCIS2 systems, including the internal optomechanics and units interconnecting these chambers. The 3D model showing details of individual systems will be provided to the Supplier after signature of the contract.

## References to standards

Whenever this document refers to standards or standardized/ standardizing technical documents the FZU allows also equivalent solution to be offered. If the Supplier offers equivalent solution the FZU shall not reject his bid once the Supplier proves, by appropriate means in the bid, that the offered supplies, services or works meet the requirements including references to standards or technical documents in an equivalent manner.

# Scope of work

Within this procurement the following scope of work is required:

- Final production design and manufacture of three (CIS1A, CIS1B and CIS2) vacuum chambers and fixings, based on the detailed conceptual design provided by FZU.

- Final production design and manufacture of four optical support chassis complete with isolation legs, based on the detailed conceptual design provided by FZU.

- Final production design of the eleven optomechanical mounts OM1, OM1.5, OM2, OM3, OAP1, OM3.5, OM3SA, OM4, OM4.5, OM5 and MBD based on the detailed conceptual design provided by FZU.

The Supplier is specifically required to:

- Develop detailed design of the individual optomechanics, based on the provided conceptual drawings

- Identify and implement suitable actuators meeting the positioning requirements

- Develop design of attachment of the optomechanics to the respective optical support chassis; the solution should ideally use brackets not extending the optomechanics’ footprint on the support chassis

- Manufacture of ten optomechanical mounts OM1, OM2, OM3, OAP1, OM3.5, OM3SA, OM4, OM4.5, OM5 and OMBD.

- Design, supply, installation and testing of the optomechanical system controller including all actuators and associated devices, according to the functional specification supplied by FZU (see Section 9.2).

- Design, supply, installation and testing of all cabling, cable management and vacuum feedthroughs for equipment mounted in and on all three CIS chambers.

- Final production design, supply and testing of the laser shutter in CIS1A, based on the detailed conceptual design provided by FZU.

- Supply of connecting vacuum tube sections, isolation bellows and other flanges and fittings as specified for CIS1A, CIS1B, and CIS2.

- Supply of Turbomolecular Pumps (TMPs), vacuum gauges, air admittance valves, and all other devices comprising the CIS vacuum system hardware (see Section 8.1).

- Design, supply and testing of the vacuum control system according to the functional specification supplied by FZU (see Section 9.1).

- Final cleaning, trial assembly and functional testing (vacuum pump-down and control functions) of the entire CIS, as far as possible at works.

- Packaging for transport and delivery of all the CIS system components to the ELI-Beamlines facility.

Installation of the CIS chambers at ELI-Beamlines and final integration of the controls is not required within this procurement package.

The supply of beam-contacted optical components (i.e. mirrors, lenses, windows, etc.) is not part of the scope of supply. The beam dump is also not part of the scope of supply. If necessary (e.g. for testing) any such components will be loaned to the Supplier by the Client (FZU).

# Arrangement of the CIS components and integration with L4 laser system

The overall arrangement of the CIS components in relation to the L4 Final Amplifier and the L4 Compressor is shown in Figure 1. A more detailed arrangement of chambers CIS1A, CIS1B and CIS2 in the building is shown in Figure 2. General arrangement drawings of each of the chambers are included in the Drawing Package at Section 13.1.1.

Chambers CIS1A and CIS1B are fixed directly to the floor of Laser Hall L4b, whereas chamber CIS2 stands on a stiff structural frame and is also secured to the wall of Compressor Room L4c.

The optical support chassis within CIS1A and CIS1B stand directly on the floor of L4b, whereas the two optical support chassis within CIS2 are each independently fixed to the wall of Compressor Room L4c.

Lengths of vacuum tube connect between the chambers with, in each case, a bellows on one end to allow for relative movement of the chambers during pump-down. Not visible in Figure 2 is a tube that passes from chamber CIS2 through a wall penetration into Experimental Hall L3.

CIS1A

CIS1B

CIS2

CIS1A

CIS2

CIS1B

Figure 2: General arrangement of the chambers CIS1A, CIS1B, and CIS2 in the building.

The beam path through the complete Compressor Image relay System is shown in Figure 3. The mirror M1 reflects the diverging beam exiting the L4 laser system, and the beam is re-collimated by the off-axis parabola OAP1. The re-collimated beam is then sent from the L4b laser hall to the floor below by mirror M4. The mirror M5 serves to reflect the beam to the 10 PW compressor.

Figure 3: Beam path through the complete Compressor Image relay System. Mirrors shown in CIS1A correspond to configuration for the S-polarization on the gratings. The translatable beam dump mirror MBD in CIS1A is not shown in this figure.

The beam path between chambers CIS1A and CIS1B contains a window that provides vacuum separation between the two chambers. The vacuum condition within chamber CIS1A is determined entirely by its connection to vacuum chamber VC2, which houses the final amplification stage of L4 laser. The tube interconnecting CIS1A and CiS1B must be installed in a very cramped location beneath a laser table and will be of square cross section (see Section 4.6)

Chambers CIS1B and CIS2 are always directly connected to each other with round tube and vacuum is established within these chambers by a dedicated vacuum system which is described in detail in Section 8.1 (CIS vacuum circuit and its connection to ELI-Beamlines central vacuum).

Chamber CIS2 is connected to the L4 Compressor chamber through an actuated isolation valve (DN630). Control of this valve is conditional on the vacuum status on either side, therefore the CIS vacuum control system must interface with the L4 Compressor control system. Further details are given in Section 9, CIS Control System.

The round tube providing the physical connection into the E3 Experimental Hall will be terminated in a blank flange. If this connection is used in future to transfer the uncompressed beam then the blank flange will be replaced by a window to ensure continued separation of the gas spaces.

Table 1 provides a summary of functional requirements for all the optomechanical mounts that form part of the CIS and shows in which chamber they are located. While the individual parameters in Table 1 are binding the Supplier is entitled to select appropriate type of vacuum-compatible electrical actuator for each mount /axis. If needed, FZU can provide Newport VHRU and/or LTA vacuum actuators but the Supplier must provide full integration of these actuators into the optomechanics and into the control system.

Table 1: CIS Optomechanics (note: the individual optomechanics OMx correspond to the respective mirrors Mx).

| Optomechanical mount No. | Range | Resolution 1 | Accuracy 2 | Actuator type | Encoder |
| --- | --- | --- | --- | --- | --- |
| Chamber CIS 1A |  |  |  |  |  |
| OM1 |  |  |  |  |  |
| Yaw | ±1° | 2 µrad | N.A. | Manual | - |
| Pitch | ±1° | 2 µrad | N.A. | Manual | - |
| OMBD (to beam dump) |  |  |  |  |  |
| Translation | 380 mm | N.A. | N.A. (mechanical limit) | Stepper motor or pneumatic | End position detection |
| OM1.5 (supply not part of this contract) |  |  |  |  |  |
| Yaw | ±1° | 2 µrad | N.A. | Manual | - |
| Pitch | ±1° | 2 µrad | N.A. | Manual | - |
| OM2 |  |  |  |  |  |
| Yaw | ±1° | 2 µrad | N.A. | Manual 3 | - |
| Pitch | ±1° | 2 µrad | N.A. | Manual 3 | - |
| Chamber CIS 1B |  |  |  |  |  |
| OM3 |  |  |  |  |  |
| Yaw | ±1° | 1µrad | 5 µrad | Stepper motor | - |
| Pitch | ±1° | 1 µrad | 5 µrad | Stepper motor | - |
| Axial translation | ±12.5 mm | 5 µm | 5 µrad | Stepper motor |  |
| OM3SA (sub-aperture) |  |  |  |  |  |
| Yaw | ±1° | 2 µrad | N.A. | Manual | - |
| Pitch | ±1° | 2 µrad | N.A. | Manual | - |
| OAP1 |  |  |  |  |  |
| Yaw | ±0.5° | 2 µrad | 5 µrad | Stepper motor | Y |
| Pitch | ±1° | 2 µrad | 5 µrad | Stepper motor | Y |
| Axial translation | ±25 mm | 5 µm | 10 µm | Stepper motor | Y |
| OM3.5 |  |  |  |  |  |
| Yaw | ±1° | 2 µrad | N.A. | Manual | - |
| Roll | ±1° | 2 µrad | N.A. | Manual | - |
| Rotation (flip) | 95° | 2 µrad | 10 µrad | Stepper motor | End positions detection |
| OM4 |  |  |  |  |  |
| Yaw | ±1° | 1 µrad | 5 µrad | Stepper motor | Y |
| Pitch | ±1° | 1 µrad | 5 µrad | Stepper motor | Y |
| Chamber CIS 2 |  |  |  |  |  |
| OM4.5 |  |  |  |  |  |
| Yaw | ±1° | 1 µrad | 5 µrad | Stepper motor | Y |
| Pitch | ±1° | 1 µrad | 5 µrad | Stepper motor | Y |
| Translation | 400 mm | N.A. | N.A. (mechanical limit) | Stepper motor | End position detection |
| OM5 |  |  |  |  |  |
| Yaw | ±1° | 1 µrad | 5 µrad | Stepper motor | Y |
| Pitch | ±1° | 1 µrad | 5 µrad | Stepper motor | Y |

1 One incremental step of motorized actuator; one graduation of fine adjustment scale for micrometric screw

2 Absolute position knowledge by encoder readout, bidirectional repeatability for motions without encoder

3 Design of the mount should enable implementation of a stepper motor in the future

Yaw: rotation around the vertical axis

Pitch: rotation around the horizontal axis parallel to the mirror surface

Roll: rotation around the horizontal axis perpendicular to the mirror surface

# Specific design of CIS1A unit

This section describes the specific design of the elements comprising CIS1A. Generic design and performance requirements that relate to all vacuum chambers and vacuum components are covered in Sections 7  and 8.

## CIS1A Optical Configuration

The beam path through CIS1A is shown in Figure 4. The section of the beam path enclosed by the default optical configuration (“S-polarization”) of unit CIS1A includes fixed mirrors M1, M2 and movable mirror MBD. The movable mirror serves to send on demand the laser beam to the beam dump attached to the corresponding flange on the CIS1A chamber.

There is an alternative optical configuration (“P-polarization”) of unit CIS1A that includes a third fixed mirror M1.5. The final design for both configurations is included in the scope of work but the extra optomechanical mount OM1.5 for the alternative (P-polarization) configuration is not to be manufactured within this contract.

Figure 4: Laser beam path through CIS1A. The fixed mirrors M1 and M2 are shown in blue, the movable mirror MBD appearing in green is shown in the retracted position for the S-polarization configuration. The laser shutter (see text further below in this document) is represented in yellow.

## CIS1A Vacuum Chamber

The CIS1A vacuum chamber is designed as an externally ribbed structure in the shape of a rectangular box with approximate overall outer dimensions 1,923 mm high, 1,265 mm wide and 1,095 mm deep (see Section 13.1.1, Drawing Package). External view of the chamber is shown in Figure 5. The chamber has a removable top lid and hinged access doors on two sides. There are several circular flanges whose details are given in the CIS1A Flange Schedule in Table 8. The chamber shall stand on four height-adjustable feet that will be secured directly to the L4b laser hall floor with anchor bolts.

The top lid includes a penetration for the operating shaft of a vertical laser shutter. Vacuum integrity is maintained by enclosing the shaft in a long-travel bellows allowing at least 400 mm movement.

Figure 5: CIS1A chamber body. The side doors and top lid appear in blue, the actuating system of the laser shutter is shown in yellow.

The chamber body shall be manufactured from stainless steel. The chamber structure has nominal wall thickness of 20 mm - the value was set by FEM analysis to limit deformations due to atmospheric pressure. The Supplier shall repeat the FEM analysis of the finally detailed design.  The top lid and side doors shall also be manufactured from stainless steel. One of the side doors includes the ISO-K (or equivalent solution) 320 mm diameter flange for mounting the water-cooled beam dump (not part of this contract). The design of this door has two configurations corresponding to the S and P-polarization options with the beam dump in different positions. Drawings for manufacture both doors shall be prepared but only the “S-polarization” configuration door shall be manufactured.

## CIS1A Optical Support Chassis

The optical components in each vacuum chamber section of the CIS are supported in optomechanical mounts, which in turn are fixed to an optical support chassis. The optical support chassis will be fixed rigidly to the building structure using isolation legs, which are mechanically decoupled from the chamber vessel by means of bellows. Thus, for CIS1A the optical support chassis will be bolted directly to the L4b hall floor with four isolation legs, independently of the vacuum chamber supports.

The optical support chassis will be unaffected by any deformation of the vacuum chamber as it is pumped down from atmospheric pressure.

The arrangement of the CIS1A chassis is shown in Figure 6.  The Supplier shall develop the conceptual design provided by FZU to the detailed design. The chassis shall be designed so that it can accommodate both the S- and P-polarization configurations without modification.

The optical support chassis and the isolation legs are part of the scope of supply as they are vacuum and mechanically integrated with the vessel and must be assembled with the vessel during factory vacuum and structural testing of the unit.

Both the optical support chassis and the isolation legs shall be manufactured from aluminium or stainless steel. The optical support chassis should be free of thermal and mechanical stress. Therefore if arc welding is used, which might introduce thermal stresses into the material, stress relieving by heat treatment should be applied afterwards. Ideally laser welding should be used. Alternatively, the optical support chassis can be bolted from several smaller units. Any bolted connections must avoid any trapped volumes of air and, e.g. mounting holes shall be tapped either right through or shall be equipped with venting bores.

Figure 6: CIS1A optical support chassis onto which the optomechanical mount will be fixed. The isolation legs are connected to ISO-K (or equivalent solution) 100 flanges on the chamber by bellows seals to provide mechanical decoupling from the chamber vessel.

Design of the isolation legs and its internal structure are shown in more detail in Figure 7. The structure consists of double bellows providing full isolation of the force on the flange used for the leg, arising from pressure difference between outer atmospheric pressure and the chamber vacuum. As seen from Figure 7, the outer atmospheric pressure acting on the flange is transferred by 3 rod-shaped independent beams to the chamber body and the force acting on the support leg is zero.

Figure 7: Concept of the CIS1A isolation legs using double bellows system. The double bellows fully remove the force arising from the atmospheric pressure – vacuum pressure differential on the flange cross-section, as the force is fully transmitted on the vessel by means of 3 cylindrical beams.

## CIS1A Optomechanical Mounts

The outline technical requirements of the optomechanical mounts for CIS1A are given in Table 1. The mirror mounts are fitted with actuators that allow a fine adjustment of yaw and/or pitch by manual actuators or by operation of stepper motor from the external control system. The beam-dump mirror mount (OMBD) must be capable of rapid sideways translation, in less than 5 seconds, over a distance of 380 mm.

The mirrors will be mounted in frames which will be installed into the supplied optomechanical mounts and must thus interface with these mounts. Dimensions of the individual CIS1A mirrors and of their mounting frames are shown below in Table 2.

|  | Substrate size |  |  | Frame size |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Mirror | Width (mm) | Length (mm) | Thickness (mm) | Width (mm) | Length (mm) | Thickness (mm) |
| M1 | 250 | 350 | 60 | 310 | 400 | 58 |
| MBD | 250 | 350 | 60 | 290 | 390 | 58 |
| M1.5 | 350 | 250 | 60 | 310 | 400 | 58 |
| M2 | 350 | 250 | 60 | 310 | 400 | 58 |

Table 2: Size of the CIS1A mirrors and of the mounting frames (manufacture of M1.5 is not part of this supply)

## CIS1A Laser Shutter

The purpose of the laser shutter is to provide on demand light tight isolation of the CIS1A chamber from the downstream optical path. The shutter must not present a vacuum barrier, in order not to create a pressure difference across the blade of the shutter.

The shutter will be operated for reasons of personal safety when it is required that no laser light propagates beyond the CIS1A environment.

The aperture in the shutter is 400 mm so that the blade stroke must exceed this value to provide an overlap.

The blade has to be made of two parts. The mechanical strength will be provided by a metal plate and the laser facing side must be equipped by a neutral density (ND) absorbing glass sheet with dimensions 280x280 mm and thickness 10 mm. The glass sheet will be supplied by FZU.

The shutter has to be actuated pneumatically. It has to be capable to fully close in less than 3 seconds, ideally in 1 second.

The pneumatic actuator will be mounted to the lid of the CIS1A chamber, as indicated in Figure 8. The Supplier shall develop the full design of the actuating mechanism of the shutter, based on the conceptual design provided by FZU. The actuator shall be equipped with full open and full closed position switches.

Figure 8: Conceptual scheme of the laser shutter (in yellow) and of the actuating pneumatic mechanism (electro-pneumatic actuator shown in purple, vacuum bellows in dark blue) integrated onto the CIS1A chamber.

## Interconnecting tube between CIS1A and CIS1B chambers

Chambers CIS1A and CIS1B shall be connected by a square vacuum tube (internal cross section 300 by 300 mm, external cross section not exceeding 340 x 340 mm) with length of 1600 mm, equipped with a bellows at one end (preferably CIS1A) and a wedged optical window at the other end (preferably CIS1B). This tube has to pass under optical table number 17 of the L4 laser (the width of the table is 1,500 mm) system where access is restricted.

The mount of the optical window (diameter 500 mm, thickness 70 mm) must provide vacuum isolation for pressure differential from either side of the window. The design shall be based on the conceptual design provided by FZU (see Figure 9 below) and shall be coordinated with the FZU.

Figure 9: Conceptual scheme of the vacuum interconnecting square tube between the CIS1A and CIS1B chambers. The external cross section must not exceed 340x340 mm. The tube is connected with CIS1A by a bellows (not represented here, will be supplied by FZU). The optical vacuum window is mounted in a segment adjoining the CIS1B chamber, see conceptual design in the right picture. The laser table above the interconnecting segment is not represented in this figure.

# Specific design of CIS1B unit

This section describes the specific design of the elements comprising the CIS1B unit. In addition, generic design and performance requirements that relate to all vacuum chambers and vacuum components are covered in Sections 7  and 8.

## CIS1B Optical Configuration

The beam path through CIS1B is shown in Figure 10 and includes fixed mirrors M3, M3SA, off-axis parabola OAP1, M4, and moveable mirror M3.5. This moveable mirror M3.5 flips on demand into the laser beam trajectory and sends the beam outside the CIS1B chamber, through rectangular optical window.

Figure 10: Beam path through CIS1B showing assembly of the optomechanical mounts of OM3, OM3SA, OAP1, OM3.5 and OM4. The nominal path of the laser beam is M3-OAP1-M4. The sub-aperture mirror M3SA, picking up a fraction of the beam transmitted through M3, reflects the sub-aperture beam through optical window outside CIS1B for diagnostic purposes. The mirror M3.5 flips into the main beam, and is capable to send the entire beam, through optical window, outside CIS1B for diagnostic purposes.

## CIS1B Vacuum Chamber

The CIS1B vacuum chamber is designed as an externally ribbed structure in the shape of a rectangular box with approximate overall outer dimensions 945 mm high, 1,844 mm wide and 2,060 mm deep (see Section 13.1.1, Drawing Package). External view of the chamber is shown in Figure 11. The top lid design can be adapted from the conceptual design developed by FZU. There are numerous flanges whose details are given in the Flange Schedule in Table 9.

Figure 11: CIS1B chamber body. The chamber interior is accessed by a lid which must be lifted in a highly space constrained environment of the L4 adjacent systems.

The chamber body and lid shall be manufactured from stainless steel. The chamber structure has nominal wall thickness of 20 mm. The Supplier shall make FEM analysis and use it for optimizing the finally detailed design.

The chamber shall stand on four height-adjustable feet that will be secured to the L4b laser hall floor with anchor bolts.

## CIS1B Optical Support Chassis

The CIS1B optical support chassis will be bolted directly to the L4b hall floor with four isolation legs, which are mechanically decoupled from the chamber vessel by means of bellows. Thus, the CIS1B the optical support chassis will be independent of the vacuum chamber and will be unaffected by any deformation of the vacuum chamber as it is pumped down from atmospheric pressure. The arrangement of the CIS1B optical support chassis is shown in Figure 12.

The table of the support chassis (140 mm thickness) shall be manufactured from one single block of cast aluminium free of internal stress. The layout of the support chassis has to take into account strong height constraints given by the position of the CIS1B mirrors above the floor level of the L4b laser hall. For this reason the isolating bellows blocks use a specific compact arrangement with respect to the optical support chassis, as shown in Figure 12.

Figure 12: CIS1 B optical support chassis configuration with the aluminium table and isolating bellows blocks.

The conceptual design of the isolating double bellows for the CIS1B chamber is shown in Figure 13. The isolating bellows will be attached to accurately positioned encast steel plates in the L4b hall floor. The Supplier shall develop detailed design of the isolating legs. The design must take into account the height limitations given by the position of the laser beam axis 400 mm above the L4b hall floor level.

Figure 13: Concept of the CIS1B isolation legs. The legs must take as minimum height as possible due to height constraints.

## CIS1B Optomechanical Mounts

The outline technical requirements of the optomechanical mounts for CIS1B are given in Table 1. All of the mirror mounts are fitted with actuators that allow fine adjustment of yaw and/or pitch by operation of stepper motor from the external control system. The mount for off-axis parabola OAP1 allows for ±25 mm axial translation and ±5 mm vertical and lateral translation and mirror mount OM3.5 for 95° rotation by operation of stepper motor from the external control system.

All CIS1B mirrors will be mounted in frames supplied by FZU. The frames will be inserted and attached to the optomechanical mounts. Dimensions of the individual CIS1B mirrors and of their mounting frames are shown below in Table 3.

|  | Substrate size |  |  | Frame size |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Mirror | Width (mm) | Length (mm) | Thickness (mm) | Width (mm) | Length (mm) | Thickness (mm) |
| M3 | 350 | 350 | 60 | 390 | 390 | 60 |
| M3SA | 152.4 | 152.4 | 25 | N.A. | N.A. | N.A. |
| OAP1 | 380 | 380 | 75 | 430 | 430 | 65 |
| M3.5 | 380 | 520 | 90 | 430 | 570 | 50 |
| M4 | 380 | 520 | 90 | 430 | 570 | 80 |

Table 3: Size of the CIS1B mirrors and of the mounting frames

# Specific design of CIS2 unit

This section describes the specific design of all the elements comprising CIS2. In addition, generic design and performance requirements that relate to all vacuum chambers and vacuum components are covered in Sections 7  and 8.

## CIS2 Optical Configuration

The beam path through CIS2 is shown in Figure 14 and includes fixed mirror M5 and moveable mirror M4.5. The mirror M4.5 can be translated horizontally by 400 mm and if inserted it deflects the downward travelling beam into Experimental Hall L3. If M4.5 is withdrawn the beam continues downwards to M5 where it is deflected into the L4 10 PW compressor.

Figure 14: Laser beam path through CIS2 showing mirrors M4.5 and M5 and the corresponding optomechanical mounts OM4.5 and OM5. In the nominal configuration the mirror M4.5 is retracted and the beam propagating vertically is reflected by the lower lying mirror M5 towards the 10 PW compressor.

## CIS2 Vacuum Chamber

The CIS2 vacuum chamber is designed as an externally ribbed structure in the shape of a rectangular box with approximate overall outer dimensions 2,550 mm high, 1,400 mm wide and 1,030 mm deep (see Section 13.1.1, Drawing Package). External view of the chamber is shown in Figure 15. The chamber has a hinged access door on its front face. There are numerous circular flanges whose details are given in the Flange Schedule in

Table 10.

The chamber body and door shall be manufactured from stainless steel. The chamber structure has nominal wall thickness of 15 mm (the structure is designed with to minimise its weight and is ribbed with ribs 80 mm tall). The Supplier shall make the FEM analysis of the finally detailed designed to optimize the rib structure resulting in deformations of the chamber walls not exceeding 1 mm.

Figure 15: CIS2 chamber body. The hinged rectangular door serves to access interior of the chamber. The steel frame helping to support the chamber is represented in yellow. The bellows on the entrance flange of the CIS2 chamber for installation at ELI-Beamlines will be provided by FZU.

The chamber is principally supported on a rigid steel frame that is secured to the wall of Compressor Room L4c. and the chamber itself is also anchored firmly to the wall. The wall is constructed from reinforced concrete into which a pattern of cast-in steel fixing channels has been integrated. Details of the fixing channels and the concrete reinforcement pattern will be provided to aid the Supplier in designing suitable attachments for the support frame and chamber.

The rear face of the CIS2 chamber is close to the L4c hall structural wall and specific considering must be given in the design with respect to installation of the chamber and of the internal optical support chassis. Also the design must consider installation of the connecting ISO 500 (or equivalent solution) vacuum tube into the wall penetration between the L4c and E3 halls; the Supplier may modify the conceptual design developed by FZU.

## CIS2 Optical Support Chassis

There are two separate optical support chassis, one for each mirror mount M4.5 and M5. Each chassis is fixed directly to the compressor room L4c wall using isolation legs to provide a vacuum seal but avoiding movement during pump down as a result of chamber deflection. Fixing considerations are similar to that described for the chamber body in section 6.2. The Supplier must perform FEM analysis of the final design of the support chassis supported by the isolating bellows and verify absence of low-frequency resonances (below and including 50 Hz).

Figure 16: CIS2 chamber interior showing arrangement of the two optical support chassis (left). The bellows structure providing isolation of the support chassis from the chamber is on the right.

## CIS2 Optomechanical Mounts

The outline technical requirements of the optomechanical mounts for CIS2 are given in Table 4. Both mirror mounts are fitted with actuators that allow a fine adjustment of the yaw and pitch by operation of stepper motor from the external control system.

|  | Substrate size |  |  | Frame size |  |  |
| --- | --- | --- | --- | --- | --- | --- |
| Mirror | Width (mm) | Length (mm) | Thickness (mm) | Width (mm) | Length (mm) | Thickness (mm) |
| M4.5 | 380 | 520 | 90 | 430 | 570 | 80 |
| M5 | 380 | 520 | 90 | 430 | 570 | 80 |

Table 4: Size of the CIS2 mirrors and of the mounting frames

The Supplier shall design a self-contained mechanical handling aid for each optomechanical mount so that it can be securely positioned outside the chamber and then installed horizontally and or vertically directly by means of guide rails into its respective support chassis position through the relevant chamber access door.

# General Requirements for Vacuum Chamber Design and Manufacture

Applicable to stainless steel components including vacuum chambers

- Physical pre-cleaning

  - Prior to cleaning, all components will be examined, especially all vacuum interface surfaces (flanges)

  - Vacuum surfaces of the flanges must be inspected one by one and must be free of scratches, rolling, cracks, scale, or other defects.

  - All swarf, burrs, etc. from the machined surfaces will be physically removed. The procedures may include high pressure air blasting, water jet, scraping, swabbing etc.

  - Gross contaminations from the manufacturing process will be removed by washing or rinsing with suitable general purpose solvent.

- Cleaning / washing

  - All vacuum components must be cleaned separately.  Cleaning of assembled components and systems is not allowed.

  - All components to be cleaned will be thoroughly inspected. Areas or re-entrant volumes which may trap cleaning agents or may have deposits from the fabrication process must be identified and closely followed during the whole cleaning process.

  - Degrease using detergent and town water, using a mild alkaline non-etch environment with a pH <11.  If local regulation permits it is advised to use as an alternative trichloroethylene or perchloroethylene. Allowed procedures include rinsing, immersion and swabbing.  Use of high pressure hot water jet at temperature 60-80° C is recommended.

  - Application of vapour degreasing process additionally to c) is recommended.

  - Wash thoroughly with town water, without detergent.  Recommended is use of high pressure hot water jet at temperatures around 80° C or higher. Pay particular attention to any trapped volumes.

  - Wash with clean demineralised water (minimum resistivity = 106 Ohm/cm). Allowed procedures include immersion and rinsing; use of high pressure hot water jet (temperature 60-80° C) is recommended.

  - Step f) should be ideally done by using ultrasonic bath. Use of ultrasonic bath for all small to medium size components and complex-shape components in step f) is mandatory. For L4 CIS this is the case for the support structures and of all optomechanical mounts that will be installed inside the chambers.

  - Ideally all components should be baked out at temperatures 200 to 300° C in a vacuum oven at a pressure <1 mbar, to achieve vacuum surfaces low in hydrocarbons.  While application of this technique for the L4 CIS chambers may be limited in terms of size, its use is mandatory for all optomechanics.

  - Vacuum baking as specified in step h) is mandatory for bellows. Extended period (>10 hours) of baking shall be applied.

Specific procedures for cleaning stainless steel bellows

- Bellows must be cleaned in extended configuration.

- Step c) from above is not to be applied. Especially no alkaline degreasing solutions must be used as these are prone to particulate deposition.

- Recommended cleaning procedure includes step g) from above: immersion in ultrasonic bath of ethanol or isopropyl.

- Cleaning should include thorough drying of the bellows using clean dry air or nitrogen and/or drying in air oven at temperature of 100-150° C, for a period of at least 1 hour.

Testing and inspection

The chamber will be vacuum tested at works to prove that the vacuum cleanliness criteria were met and that it is capable of reaching and maintaining the required 10-7mbar vacuum.  This demonstration will require the fitting of the quadrupole mass spectrometer, the TMPs and some of the vacuum gauges that will be used in the final assembly at ELI.

Packaging

The external surfaces of the CIS chambers will be clean to a level commensurate with operation in a cleanrooms of Class 5 according to ČSN EN ISO 14644 standard (equivalent to ISO 14644).  The external surface may be painted, except on sealing faces. To preserve this level of cleanliness during shipping and installation a multi-layer wrapping protocol will be applied.

The vessel will be supplied with doors or covers fitted to all apertures. The chamber will be filled with clean dry air or nitrogen.

The CIS chambers will be triple bagged, with two layers of polyethylene cleanroom film (e.g. UltraLOPlus™ or equivalent) and then one layer of a vapour barrier material such as foil.  Seams will be taped or heat sealed.  Excess air between wrappings will be evacuated and the bags sealed with desiccant packs inside.

General rules for assembling

- Any use of grease shall be avoided.

- O-rings will be stored in a sealed plastic bag. Any use of solvents in cleaning O-rings shall be avoided. The allowed technique for cleaning O-rings from hydrocarbons is baking in vacuum oven at temperatures <180°C.

- Prior to installation of any O-ring the corresponding flange seal surfaces must be cleaned with alcohol (ethanol or isopropyl) immediately before installation.

# General Requirements for Vacuum Equipment Design and Supply

As a working principle all operations carried out within any CIS chamber at ELI-Beamlines will be consistent with Class 100 operations. It is envisaged that the relevant chamber will be enclosed in a modular Class 100 containment for initial installation at ELI-Beamlines and whenever the chamber is opened for major work. Sufficient space will be provided within this modular enclosure to remove and handle the largest components. Dry compressed air (ISO 8573 class [1,1,1] available at ELI-Beamlines facility) bleed may be used to avoid mixing the chamber environment with room air.

All exposed metal parts (chamber, tubing, valves, etc.) will be electrically bonded to the facility ground for machine and personnel safety. The design of the chamber and its major components shall provide where necessary tapped holes for attachment of bonding straps.

The outer surface of the chamber and/or of the doors shall be vinyl film wrapped or locally painted. Details will be agreed between the Supplier and FZU during the early manufacturing phase.

## CIS vacuum circuit and its connection to ELI-Beamlines central vacuum

All the three CIS chambers being subject of this RSD are designed to provide vacuum conditions of 10-7 mbar.

Pumping of the CIS1A chamber will be accomplished through its connection to the vacuum circuit of the L4b laser, located in the L4b hall. As the CIS1A chamber is vacuum separated from the other two CIS chambers (and from the L4 10PW compressor) by means of the vacuum optical window there are no CIS1A vacuum elements controlled by the CIS vacuum controller.

The CIS1B and CIS2 chambers will be pumped to roughing vacuum (10-2 mbar) by using one dedicated branch of the central primary vacuum system of the ELI-Beamlines facility. This consists of stainless steel tubing (ISO-K 160) which emerges from the floor channel under the CIS2 chamber, see Figure 17. The roughing pumping is made using an ISO-K 100 flange on the underside of the CIS2 chamber (see also Section 10, Vacuum flange schedule). A separate L4c dedicated circuit (ISO-K 100) is provided for maintaining backing vacuum (10-2 mbar) for turbomolecular pumps (TMP) used to maintain high vacuum in the chambers. Both primary vacuum circuits are pumped by dry pumps.

Figure 17: Schematics of the primary vacuum (roughing and backing circuits) supply for the CIS system in the L4c hall. Design of the backing line layout providing vacuum to TMP on each of the CIS1B and CIS2 chamber is represented only schematically.

All CIS1A, CIS1B and CIS2 chambers will use vacuum flanges of the ISO-K standard or equivalent solution for HV, sealed by  fluoroelastomer O-rings. All CIS1A flanges are sealed by single (standard) O-rings on all vacuum ports. The large lid of CIS1B and the access door of CIS2 will use a double O-ring seal arrangement to minimize overall permeation through the corresponding seals. While the conceptual design developed by ELI-Beamlines fully takes into account the double O-ring arrangement, its detailed design shall be elaborated by the Supplier.

A block scheme of the vacuum circuit of the CIS chambers is shown in Figure 18. All elements indicated in this scheme, except the large valves ISO-K (or equivalent solution) 630 and ISO-K (or equivalent solution) 400, are included in the scope of the contract. The vacuum scheme also involves the associated segment of the ELI-Beamlines primary vacuum, which is supplied and which will be controlled by the L4c hall ELI-Beamlines facility controls.

Figure 18: Block scheme of the CIS vacuum circuit. The CIS1A chamber is connected to the L4 laser vacuum system and isolated by a window from CIS1B and CIS2 chambers. The elements in the Plant Room will be supplied by FZU. The lid of CIS1B and the door of CIS2 chambers are equipped with double O-rings. The large valves No. 10 and No. 19 are not required within the scope of this supply although the corresponding channels in the control system are required to be supplied. Also the MS detector (No.15) is not part of the supply.

In the operation the CIS1B and CIS2 chambers will have to be fully isolated from the adjacent vacuum systems by valves. The valve #10 between CIS2 and the 10 PW compressor and also valve #19 between CIS2 and downstream beam transport elements are important elements of the CIS machine safety.

The CIS1B and CIS2 chambers will be equipped by a set of vacuum gauges (#11, #12, #13, and #14). The sensors will involve Bayard-Alpert ionization gauge due to its inherent precision at low pressures, and atmospheric gauges, for machine safety, providing accurate measurements of pressure when the chambers are vented by clean air.

The vacuum cleanliness will be on-line monitored by a COTS compact quadrupole mass spectrometer, providing the partial pressures of vacuum residuals including water vapour, hydrocarbons and volatile organic compounds. This mass spectrometer will be supplied by FZU. Readout of this spectrometer will be an important element of compressor machine safety.

## Cable trays

The Supplier shall develop the design of both external and internal cable trays.

The CIS1B and CIS2 units will be controlled from the control rack(s) located in the L4c hall. The CIS1B external cable trays must therefore make it possible leading the electrical cables to the L4c hall to the vacuum and optomechanical controllers. Exact position of these controllers in the L4c hall will be agreed between FZU and Supplier during early phase of this project.

The internal (vacuum) cable trays shall serve to arrange cables from the motorized actuators of the individual optomechanics and other electrical devices within the chambers to the vacuum feedthroughs (see Section10, Vacuum Flange Schedule). The cable trays must follow accessible routes and should preferentially be attached to the vacuum chamber body; if necessary these trays can also be attached along the optical support chassis and/or its support structure.

The internal cables and trays must be made entirely from high-vacuum compatible materials and their design must avoid any cavities and/or trapped volumes of air.

External cable trays shall be designed to route and support cables from the feedthroughs and from externally mounted equipment in an orderly fashion. The routings shall avoid clashes with doors, covers and flanges that need to be free to be opened or removed.

# CIS Control Systems

## Vacuum Control System

The CIS Vacuum Control System (VCS) must perform the following key functions: monitoring of the vacuum system condition and detection of faults, automation of vacuum processes such as pump down and venting, and safety functions related to the two output gate valves and a laser safety shutter and laser beam dump.

The VCS must be designed to the specified state machine scheme so that the state of the system is deterministic at all times. The VCS must interface to the external machine safety systems (MSS) and be designed to comply reliably with the permissive signals provided and reliably output status and fault signals to this system as specified. The VCS must also export process variables to the L4 laser SCADA interface using the Real-Time Server (RTS) of software libraries that will be provided by the Client.

The Supplier must provide a basic demonstration graphical user interface (GUI), LabVIEW-based (for reasons of compatibility with the overall L4 control system based on LabVIEW), displaying chamber vacuum system status and allowing operators to instigate sequence transitions and view pressure graphs during evacuation and venting operations. The GUI must run on any standard Windows (or equivalent solution) or Linux (or equivalent solution) PC connected by Ethernet only to the RTS.

The VCS must provide a safe, automatic response to fault events such as the failure of a single active device, vacuum leaks, vacuum contamination, loss of electrical power or air pressure, and unexpected removal of an MSS permissive input. It must communicate such failures to the external MSS via the appropriate interface signals.

The output vacuum gate valves and laser shutter are protective safety devices. Therefore, special consideration must be given to the control and position detection of these valves. All these devices are designed to close via gravity and spring-return on loss of power or pressure. Safe enabling permissives from two client-provided personnel safety systems and up to three client-provided machine safety systems must be incorporated. Removal of any permissive must 1) reliably prevent opening of the valve, and 2) cause the valve to close if open. Diagnostic coverage sufficient to detect a single failure of any component of the shutter control system that would result in a loss of these two functions must be provided. The documentation provided contains a conceptual design satisfying these requirements which the Supplier may refer to. The gate valve and shutter control system should be designed to a schema consistent with achieving safety integrity level SIL 3 according to IEC 61508, which is possible by reference to the conceptual design.

The VCS must provide a safe lockout for the chamber. This should be implemented using a lockout switch with key provided on the front panel of the control cabinet. Because of the potential to damage TMPs by removing power, the VCS must implement a controlled stop function that waits for the TMPs to signal they have stopped rotating before indicating via a pilot light that the system is stopped.

The VCS must be completely integrated using a RTS of type RMC-8354 (or equivalent solution) and programmed in LabVIEW (for reasons of compatibility with the overall L4 control system based on LabVIEW) for real-time targets. The RTS, a DIN-mount Ethernet switch (or equivalent solution), software implementation templates for the RTS and for the FPGA controllers, and 1-2 days’ training on the use of these shall be provided by the Client on the Suppliers request and delivered to the Supplier no more than 3 months after the start of the project.

Vacuum device interface shall follow a low-level implementation using one or more RIO FPGA Controllers (RIOs), selected from the Approved Devices list and employed in a similar fashion to standard industrial PLCs. These RIOs must implement state machine based control that limits the permissible machine actions according to the current state and provides sequencing of vacuum processes. Device interface via serial and Ethernet may be integrated on the RTS. However, all active devices should have a low-level interface to a RIO in addition to any high-level interface used. Disconnection of RIO devices from the RTS or from Ethernet must not interrupt any safety functionality, state machine compliance or fault detection capability.

The RIO(s) must be closely integrated with machine protection functions and must perform consistency checks on all vacuum states, valve positions, and pressure readings. RIOs must also react in case of vacuum failure via the low-level signal interface to the external MSS which will shut down the laser. This interface must be based on 24V dry-contact signals and is specified in the documentation provided. The state machine and sequences, implemented on the RIO(s), shall be regulated by the granting of the appropriate permissions by the MSS and shall be reported by signalising the appropriate state to the MSS. Devices that are not allowed to operate in particular states must be reliably prevented from activating. Unexpected removal of a permission is an external interlock and the VCS shall rapidly and autonomously act accordingly (usually by sealing off the chamber and going to a safe state). General output signal “Vacuum System Good” must further be used to signalise any detected critical alarms.

All vacuum control electronics represented between the blue dashed and red dashed lines must be mounted in one or more control cabinets depth 0.4m and height 2m located along the North wall of the L4c laser hall. Cabinets must be fitted with a plinth of 200mm for cable routing along the base and to allow integration with client’s adjacent electrical cabinets of matching dimensions.  All cabling between the chambers and the cabinet must be via floor channel and/or raised floor (provided by the Client at the time of installation) and via metal cable trays on the CIS vessels which must be provided by the Supplier. All sections of cable and pipe of length exceeding 50 cm must be supported or routed in closed metal cable trays. Cables and wires must have an appropriate jacket colour – black, grey or white for general signals, mid- to dark-blue for MSS signals, yellow or orange for personnel safety signals, and other colours for earth, mains voltages and neutral per the usual standards. Cables external to the control cabinet should be shielded to a minimum of 93% coverage and should be routed in closed metal cable trays. Special attention should be given to avoid ground loops in shield and cable tray connections.

The VCS solution should comply with the following diagrams. They show the required control system architecture and required control points (as a minimum – the Supplier may add control points so as to meet all other system requirements). Other implementation requirements are given in text notes in the diagram. Many control system hardware items in the diagram are required to be chosen from the Approved Devices list. If a device name in this diagram corresponds to an item in the Approved Devices list (Table 5), only the proven devices or an equivalent alternative explicitly approved by the Client shall be used to provide its functionality. The exact number of RIOs, terminals, interfaces, Ethernet ports and serial interfaces used is not important as long as all minimum and any additional control points are fully integrated and all other requirements are satisfied.

The diagrams contain excerpts denoted by light blue boxes containing suggested wiring, panel layout or other implementation details, and accompanying water and pneumatic diagrams are also provided. These diagrams may be used for guidance, but the precise implementation is left to the discretion of the Supplier.

Figure 19: State diagram for compliance of the VCS with the CIS concept of operations. Primary states are denoted by circles, allowed transitions by arrows and dashed lines represent permissives from safety systems. Note that both CIS chambers form the same vacuum volume (same state machine for both).

Figure 20: VCS architecture and implementation compliance diagram to be followed by the Supplier. Figure continued on the following pages.

Figure 20 continued from previous page

Figure 20 continued from previous page

Table 5: Approved devices list (vacuum controller)

Approved devices are those for which compatibility and operation with the other laser equipment is proven. The necessity for a specific implementation is compatibility with the rest of the L4 laser control systems, compatibility with the control systems of other major laser equipment of ELI-Beamlines, and for maintainability assurance. Choosing the proven implementation ensures that:

- Device has been tested by the client and is known to be proven for its intended purpose and the requirements, thus reducing development risks and ensuring operational reliability;

- Device is already used in other parts of the L4 laser, minimizing risks of integration with the L4 laser control system and ensuring that spare parts and know-how are available at the client;

- Device is known to be compatible with the client’s systems and to satisfy the client’s interface requirements, reducing the integration risks to ELI-Beamlines facility.

Device types not mentioned in this list may be freely selected by the Supplier as necessary to satisfy all other requirements.

If the Supplier does not wish to use the listed Proven Implementation they may request to use an Approved Alternative. To be approved, an alternative must have equivalent technical specifications or the Supplier must demonstrate that its specification is sufficient to satisfy all system requirements. The approved devices list should be updated and exchanged regularly as a record of all approval decisions.

| Device type | Proven implementations | Other acceptable implementations |
| --- | --- | --- |
| FPGA controller | NI 9149, NI 9147 | NI 9144, NI cRIO-9066, any approved alternative |
| FPGA analogue input | NI 9239, NI 9220, NI 9234 | Any approved alternative |
| FPGA RTD input | 4-wire: NI 9216 | Any approved alternative |
| FPGA digital I/O | NI 9423, NI 9474, NI 9477, NI 9425 | Any approved alternative |
| Other FPGA I/O | Any NI-listed C-series module | Any approved alternative |
| Passive Manometers | Pfeiffer 420C-MAN016 | Any approved alternative |
| Air pressure transducer | Festo SDE5, BD DMK331 | Any approved alternative |
| TMP >= 2300 l/s | Edwards STP-iXA3306C | Any approved alternative |
| RGA | Pfeiffer PrismaPlus | Any approved alternative |
| Control Cabinet | Rittal TS-8 | Any approved alternative |
| Serial device server | Moxa IA5450A | Any approved alternative |
| Compressed air pressure switch | Festo SDE5 with limits | SMC equivalent, any approved alternative |
| Safety-rated pneumatic control valve | Festo VOFA-L26-T32C-M-G14-1C1-APP | Any approved alternative |
| General pneumatic control manifold | Festo VTUG series | SMC equivalent, any approved alternative |

## Motion Control System

The CIS Motion Control System (MCTR) is dedicated to precision closed-loop motion control. In order to make the motion controls fully compatible with the control system of the L4 laser the low-level integration controls shall be based on the RIO platform that leverages FPGA-based feedback and distributed control architecture.

RIO modules are available with comprehensive software development kits to interface, e.g., absolute position encoders via the BiSS open source protocol. Motor interface shall utilise universal step-direction control to allow stepper power drives to be easily exchanged for alternatives if required. Due to interlock requirements, the stepper drives cannot use microstepping to increase static positioning precision, as a continuous holding current cannot be assumed (i.e., whole-step movements must be made). Ultra-high precision geared stepper should therefore be used to reach the specified resolution.

The MCTR must be designed to the specified state machine scheme so that the state of the system is deterministic at all times. The MCTR must interface to the external machine safety system (MSS) and be designed to comply reliably with the permissive signals provided and reliably output status and fault signals to this system as specified. The MCTR must also export process variables to the L4 laser SCADA interface using the Real-Time Server (RTS) and software libraries that will be provided by the Client.

The Supplier must provide a basic demonstration graphical user interface (GUI), LabVIEW-based, displaying actuator and system states and allowing operators to move axes for alignment when in the appropriate state.

The MCTR must be completely integrated using a RTS of type RMC-8354 and programmed in LabVIEW (for reasons of compatibility with the overall L4 control system based on LabVIEW) for real-time targets. The RTS, a DIN-mount Ethernet switch, software implementation templates for the RTS and for the FPGA controllers, and 1-2 days’ training on the use of these shall be provided by the Client on the Suppliers request and delivered to the Supplier no more than 3 months after the start of the contract.

The RIO(s) must be closely integrated with machine protection functions and must perform consistency checks on all drive faults and temperature thresholds for the in-vacuum RTDs. RIOs must also react to alarms and faults via the low-level signal interface to the external MSS which will shut down the laser. This interface must be based on 24V dry-contact signals and is specified in the documentation provided. The state machine and sequences, implemented on the RIO(s), shall be regulated by the granting of the appropriate permissions by the MSS and shall be reported by signalising the appropriate state to the MSS. Devices that are not allowed to operate in particular states must be reliably prevented from activating. Unexpected removal of a permissive is an external interlock and the MCTR shall rapidly and autonomously act accordingly (for example by disconnecting the drive enable input to remove current from drives). General output signal “CIS Motion Good” must further be used to signalise any detected critical alarms.

All motion control electronics represented between the blue dashed and red dashed lines must be mounted in one or more control cabinets of depth 0.4m and height 2m (Rittal TS-8, for reasons of compatibility with the cabinets used in the laser halls of ELI-Beamlines) located along the North wall of the L4c laser hall. Cabinets must be fitted with a plinth of 200mm for cable routing along the base and to allow integration with client’s adjacent electrical cabinets of matching dimensions.  All cabling between the chambers and the cabinet must be via floor channel and/or raised floor (provided by the Client at the time of installation) and via metal cable trays on the chamber vessels which must be provided by the Supplier. All sections of cable and pipe of length exceeding 50 cm must be supported or routed in closed metal cable trays. Cables and wires must have an appropriate jacket colour – black, grey or white for general signals, mid- to dark-blue for MSS signals, yellow or orange for personnel safety signals, and other colours for earth, mains voltages and neutral per the usual standards.

Special consideration must be given to the influence of electromagnetic pulse (EMP) energy on all elements of the motion control system. This is primarily ensured by providing sufficient electrical shielding of the in-vacuum components. Wherever feasible, the solution should not preclude galvanic isolation and/or fast-acting overvoltage surge protection (e.g., TVS diodes) for all connections to devices mounted in or on the chambers. Cables external to the control cabinet should be shielded to a minimum of 93% coverage and should be routed in closed metal cable trays. Special attention must be given to avoid ground loops in shield and cable tray connections.

The MCTR solution must comply with the following diagrams. They show the required control system architecture and required control points (as a minimum – the Supplier may add control points so as to meet all other system requirements). Other implementation requirements are given in text notes in the diagram. Many control system hardware items in the diagram are required to be chosen from the Approved Devices list (Table 6), for reasons of compatibility with the devices used in the L4 laser control system as well with the devices used in other major laser equipment of ELI-Beamlines. If a device name in this diagram corresponds to an item in the Approved Devices list, only the proven devices or an equivalent alternative explicitly approved by the Client shall be used to provide its functionality. The exact number of RIOs, terminals, interfaces and Ethernet ports used is not important as long as all minimum and any additional control points are fully integrated and all other requirements are satisfied.

The diagrams contain excerpts denoted by light blue boxes containing suggested wiring, panel layout or other implementation details. These diagrams may be used for guidance, but the precise implementation is left to the discretion of the Supplier as long as the implied functionality is implemented.

Figure 21: State diagram for compliance of the MCTR with the CIS concept of operations. Primary states are denoted by circles, allowed transitions by arrows and dashed lines represent permissives from safety systems.

Figure 22: MCTR architecture and implementation compliance diagram to be followed by the Supplier. Figure continued on the following pages

Figure 22 continued from previous page

Table 6: Approved devices list (motion controller)

Approved devices are those for which compatibility and operation with the existing laser equipment is proven. The necessity for a specific implementation is compatibility with the rest of the L4 laser control systems, compatibility with the control systems of other major laser equipment of ELI-Beamlines, and for maintainability assurance. Choosing the proven implementation ensures that:

- Device has been tested by the client and is known to be proven for its intended purpose and the requirements, thus reducing development risks and ensuring operational reliability;

- Device is already used in other parts of the L4 laser, minimizing risks of integration with the L4 laser control system and ensuring that spare parts and know-how are available at the client;

- Device is known to be compatible with the client’s systems and to satisfy the client’s interface requirements, reducing the integration risks to ELI-Beamlines facility.

Device types not mentioned in this list may be freely selected by the Supplier as necessary to satisfy all other requirements.

If the supplier does not wish to use the listed Proven Implementation they may request to use an Approved Alternative. To be approved, an alternative must have equivalent technical specifications or the Supplier must demonstrate that its specification is sufficient to satisfy all system requirements. The approved devices list should be updated and exchanged regularly as a record of all approval decisions.

| Device type | Proven implementations | Other acceptable implementations |
| --- | --- | --- |
| FPGA controller | NI 9149, NI 9147 | NI 9144, NI cRIO-9066, any approved alternative |
| FPGA BiSS input | SEA 9521 | Any approved alternative |
| FPGA analogue input | NI 9239, NI 9220, NI 9234 | Any approved alternative |
| FPGA RTD input | 4-wire: NI 9216 | Any approved alternative |
| FPGA digital I/O | NI 9423, NI 9474, NI 9477, NI 9476, NI 9425, NI 9426 | Any approved alternative |
| FPGA motor interface | SISU-1004, Digital I/O modules | Any approved alternative |
| Other FPGA I/O | Any NI-listed C-series module | Any approved alternative |
| Stepper motor (high load) | Newport VHRU-10V6 | Any approved alternative |
| Stepper motor (for custom integration) | Phytron VSS series | Any approved alternative |
| Stepper motor (small, low load, precision) | Newport LTAHL-PPV6, TRA-PPV6, NSA12V6 | Any approved alternative |
| Stepper motor (rotation) | Newport RVU series | Any approved alternative |
| Absolute position encoder | Renishaw Resolute series | Any approved alternative |
| Stepper motor power stage | Phytron MCD+ series; other Phytron drives | Any approved alternative |
| Serial device server | Moxa IA5450A | Any approved alternative |
| Precision limit switch | Metrol GN | Any approved alternative |
| Protection limit switch | Allectra 363 or similar construction | Any approved alternative |
| Safety limit switch | Pilz PSEN cs2.2p connected directly to SIS system | Any approved alternative |
| In-vacuum RTD | Allectra 343-PT100-C1 | Any approved alternative |

# Vacuum flange schedule

All flanges of the CIS chambers are vacuum sealed by using a fluoroelastomer O-ring seal. The top lid of the CIS1B chamber and the access side door of the CIS2 chamber shall be equipped with double Viton fluoroelastomer O-ring seal arrangement minimizing effective gas load due to O-ring permeation, with the space between the O-rings actively pumped to ~10-2 mbar by the backing vacuum circuit. Arrangement of the tubes serving to active pumping of the space between the two O-rings shall be optimized by the Supplier during preparation of manufacturing drawings.

The circular flanges shall be according to the ISO-K (or other equivalent solution for HV) and ISO-F (or other equivalent solution for HV) flange standards.

The alignment of the flanges on the body of the chambers is in some places critical and elsewhere not critical. Table 7 gives the permitted tolerance for individual tolerance grades applicable to all circular flanges specified in the flange schedule of all the three chambers CIS1A, CIS1B and CIS2.

| Flange Tolerance Grade | Tolerance Specification |
| --- | --- |
| A | Co-axial tolerance of ±1 mm or better and angular tolerance of ±0.5 degrees or better with respect to their ideal axis |
| B | Co-axial tolerance of ±3 mm or better and angular tolerance of ±2.0 degrees or better with respect to their ideal axis |

Table 7: Circular Flange Tolerance Grade

## CIS1A chamber

The chamber will be made from stainless steel grade 304L, with nominal thickness of 20 mm. The shape of the chamber and ribs provide high structural stiffness of the chamber with respect to atmospheric pressure.

The chamber will be accessed by the side hinged doors 1AF1 and 1AF2, and by the lid 1AF3. Both doors and the lid will be equipped with lifting features to be able their manipulating by the overhead crane.

The hinges of both doors will be of adjustable single or double pivot arrangement to avoid crushing the O-rings when closing the door. The doors will be aligned by locating pins providing a guide for closing. The doors must be capable of configuring the hinges on either side.

The circular flanges on the CIS1A unit are of standard type ISO-K (or other equivalent solution for HV) and ISO-F (or other equivalent solution for HV).

Classification of the CIS1A flanges is in Table 8. Note that the flange 1AF9 on the 1AF1 door must be manufactured for the “S-polarization” configuration of the laser beam (see Section 4.2).

Figure 20: Axonometric views of the CIS1A chamber showing the individual flanges, doors and lids.

| Flange | Size | Position / Purpose | Note | Tolerance |
| --- | --- | --- | --- | --- |
| 1AF1 | 800 (w) x 1110 (h) mm outer size, providing clear opening 700 (w) x 1000 (h) mm, 15 mm thickness, ribbed, side-hinged | Main access door to interior of CIS1A, access to the optics and optomechanical mounts Equipped with ISO-K 320 flange for beam dump | Fitted with lifting eyes for initial installation by overhead crane | N.A. |
| 1AF2 | 600 (w) x 1110 (h) mm outer size, providing clear opening 500 (w) x 1000 (h) mm, 15 mm thickness, ribbed, side-hinged | Access door to interior of CIS1A, access to the laser shutter | Fitted with lifting eyes for initial installation by overhead crane | N.A. |
| 1AF3 | 880 x 1060 mm outer size, providing clear opening 720 x 900 mm, 20 mm thickness, ribbed | Top lid for access to CIS1A interior, equipped with ISO-K 100 flange for the laser shutter | Fitted with lifting eyes for installation by overhead crane | N.A. |
| 1AF4 | 880 x 1060 mm outer size, providing clear opening 720 x 900 mm, 20 mm thickness, ribbed | Bottom lid of the CIS1A chamber, for easing installation of the isolation bellows | Will be attached to the encast plates in the floor of the L4b hall | N.A. |
| 1AF5 | ISO-F 100 | On the bottom lid underside, for connection of double bellows isolation of the internal chassis | The isolation bellows must be test assembled with the internal chassis at the Supplier | A |
| 1AF6 | ISO-F 100 | Idem 1AF5 | Idem 1AF5 | A |
| 1AF7 | ISO-F 100 | Idem 1AF5 | Idem 1AF5 | A |
| 1AF8 | ISO-F 100 | Idem 1AF5 | Idem 1AF5 | A |
| 1AF9 | ISO-K 320 | On the door 1AF1, for connecting the L4 laser beam dump | Initially closed by blank flange | B |
| 1AF10 | ISO-K 400 | Above 1AF2, for allowing pass of the laser beam if CIS1A mirrors are not deployed | Initially closed by blank flange | B |
| 1AF11 | ISO-K 100 | On the top lid 1AF3, for connecting bellows of the laser shutter | Equipped with the bellows and laser shutter at the Supplier | A |
| 1AF12 | ISO-K 400 | On the opposite side of 1AF1, in the bottom part of the CIS1A chamber | For connecting the interconnecting tube between CIS1A and CIS1B | A |
| 1AF13 | ISO-K 160 | Above 1AF12, contingency flange e.g. for electrical feedthrough(s) | Initially closed by blank flange | B |
| 1AF14 | ISO-K 400 | Facing 1AF10 on the other side of the chamber, for connection to the L4 laser system | Initially closed by blank flange | A |

Table 8: Positions and specifications of CIS1A vacuum flanges. Equivalent solution to ISO-K and ISO-F flange types, compatible with HV, can be provided.

## CIS1B chamber

The chamber will be made from stainless steel grade 304L, with nominal thickness of 20 mm.

The chamber interior will be accessed by the top lid 1BF1. This lid must be equipped with lifting features to be capable of lifting by a crane. The lid will be aligned by locating pins providing a guide for closing. The Supplier is allowed to modify the design of this lid developed by FZU, provided that it is capable of installation in the highly cramped surroundings. Detail will be discussed with the Supplier during the early phase of this contract.

Figure 21: Axonometric views of the CIS1B chamber showing the individual flanges, doors and lids.

| Flange | Size | Position / Purpose | Note | Tolerance |
| --- | --- | --- | --- | --- |
| 1BF1 | 1980 x 1110 mm outer size, providing clear opening 1820 x 950 mm, 20 mm thickness, ribbed | Top lid for access to CIS1B interior and for installation of the internal optical support chassis and optomechanics | Sealed by using a double fluoroelastomer O-ring seal arrangement. The ribs are bored to provide lifting eyes for installation by overhead crane. | N.A. |
| 1BF2 | ISO-F 430 (particular size outside common ISO vacuum flange standard) | On the side facing CIS1A, for installation of the vacuum window segment on the tube interconnecting CIS1A and CIS1B. | The vacuum window segment and the CIS1A-CIS1B interconnecting tube must be test installed at the Supplier. | A |
| 1BF3 | ISO-F 500 | On CIS1B underside, for connecting the vacuum tube interconnecting CIS1B and CIS2. | The flange is equipped with reinforcing ribs to prevent its displacement after pump down. | A |
| 1BF4 | Square flange 400 x 400 mm clear opening | On the North side of CIS1B, for extraction of the laser beam if reflected by the mirror M3.5. | Will be equipped with optical window, initially closed by blank flange. | A |
| 1BF5 | ISO-F 100 | On chamber underside, for connection of double bellows isolation of the internal optical table | The isolation bellows must be test assembled with the internal optical support chassis at the Supplier | A |
| 1BF6 | ISO-F 100 | Idem 1BF5 | Idem 1BF5 | A |
| 1BF7 | ISO-F 100 | Idem 1BF5 | Idem 1BF5 | A |
| 1BF8 | ISO-F 100 | Idem 1BF5 | Idem 1BF5 | A |
| 1BF9 | ISO-K 200 | On the North side of CIS1B, for electrical cable feedthroughs or contingency. | Will be equipped with cable feedthroughs if featured by the design | B |
| 1BF10 | ISO-K 200 | On the North side of CIS1B, for extraction of sub-aperture diagnostic beam reflected by mirror M3SA | Will be equipped with optical window, initially closed by blank flange. | B |
| 1BF11 | ISO-K 100 | On the North side of CIS1B, for electrical cable feedthroughs | Will be equipped with cable feedthroughs if featured by the design | B |
| 1BF12 | ISO-K 100 | On the North side of CIS1B, for electrical cable feedthroughs | Will be equipped with cable feedthroughs if featured by the design | B |
| 1BF13 | ISO-K 63 | On the South side of CIS1B, for vacuum gauge | Will be equipped with vacuum gauge | B |
| 1BF14 | ISO-K 63 | On the South side of CIS1B, for vacuum gauge | Will be equipped with vacuum gauge | B |
| 1BF15 | ISO-F 200 | On the South side of CIS1B, for TMP (2,300 l/s) | TMP mounted for the test vacuum pumping | B |
| 1BF16 | ISO-K 160 | On the South side of CIS1B, for electrical cable feedthroughs | Will be equipped with cable feedthroughs | B |
| 1BF17 | ISO-K 160 | On the South side of CIS1B, for electrical cable feedthroughs | Will be equipped with cable feedthroughs | B |
| 1BF18 | ISO-K 100 | On the South side of CIS1B, contingency | Will be blanked | B |
| 1BF19 | ISO-K 63 | On the South side of CIS1B, contingency | Will be blanked | B |
| 1BF20 | ISO-K 63 | On the South side of CIS1B, contingency | Will be blanked | B |

Table 9: Positions and specifications of CIS1B vacuum flanges. Equivalent solution to ISO-K and ISO-F flange types, compatible with HV, can be provided.

## CIS2 chamber

The chamber will be made from stainless steel grade 304L, with nominal thickness of 15 mm.

The chamber interior will be accessed by the side hinged doors 2F1. The door will be aligned by locating pins providing a guide for closing. The door will be equipped by lifting pins for manipulation by the overhead crane during installation.

Detailed design of the flange 2F6 can be modified in order to allow easy installation of the ISO 500 (or equivalent solution) vacuum tube for routing the uncompressed laser beam through the wall penetration to the E3 hall. Also, flanges 2F7 to 2F10 and 2F11 to 2F14 can be modified by the Supplier in order to optimize installation of the isolating bellows blocks.

Figure 21: Axonometric views of the CIS2 chamber showing the individual flanges, doors and lids.

| Flange | Size | Position / Purpose | Note | Tolerance |
| --- | --- | --- | --- | --- |
| 2F1 | 1400 (w) x 2440 (h) mm outer size, providing clear opening 1210 (w) x 2250 (h) mm, 15 mm thickness, ribbed, side-hinged | Access door to interior of CIS2 Equipped with two ISO-K 630 large flanges | Sealed by using a double fluoroelastomer O-ring seal arrangement. Fitted with lifting eyes for initial installation by overhead crane | N.A. |
| 2F2 | ISO-K 500 | On the top of the CIS2 chamber, for the incoming laser beam / interconnection with CIS1B. | Will connect to a bellows block mounted at the end of the vacuum tube connecting to CIS1B. | A |
| 2F3 | ISO-K 630 | On the lower part of the 2F1 flange, for the output laser beam and interconnection to the downstream 10 PW compressor. | Initially closed by blank flange. | A |
| 2F4 | ISO-K 630 | On the upper part of the 2F1 flange, for M4.5 mirror viewing purposes and/or future upgrades | Will be blanked | B |
| 2F5 | ISO-K 500 | On the underside of the CIS2 chamber, for future upgrades of the L4 system. | Will be blanked | B |
| 2F6 | ISO-K 500 | On the wall-facing side of the CIS2 chamber, for sending the laser beam to the E3 hall. | Will be blanked | A |
| 2F7 | ISO-K 80 | On the wall-facing side of the CIS2 chamber, for connection of double bellows isolation of the upper optical support chassis | The isolation bellows must be test assembled with the internal optical support chassis at the Supplier | A |
| 2F8 | ISO-K 80 | Idem 2F7 | Idem 2F7 | A |
| 2F9 | ISO-K 80 | Idem 2F7 | Idem 2F7 | A |
| 2F10 | ISO-K 80 | Idem 2F7 | Idem 2F7 | A |
| 2F11 | ISO-K 80 | On the wall-facing side of the CIS2 chamber, for connection of double bellows isolation of the lower optical support chassis | The isolation bellows must be test assembled with the internal optical support chassis at the Supplier | A |
| 2F12 | ISO-K 80 | Idem 2F11 | Idem 2F11 | A |
| 2F13 | ISO-K 80 | Idem 2F11 | Idem 2F11 | A |
| 2F14 | ISO-K 80 | Idem 2F11 | Idem 2F11 | A |
| 2F15 | ISO-K 160 | On the upper part of the South side of CIS2, for electrical cable feedthroughs | Will be equipped with cable feedthroughs | B |
| 2F16 | ISO-K 160 | On the lower part of the South side of CIS2, for electrical cable feedthroughs | Will be equipped with cable feedthroughs | B |
| 2F17 | ISO-K 63 | On the North side of CIS2, for vacuum gauge(s) | Will be equipped by the vacuum gauge(s) | B |
| 2F18 | ISO-K 63 | On the North side of CIS2, for vacuum gauge(s) | Will be equipped by the vacuum gauge(s) | B |
| 2F19 | ISO-K 63 | On the North side of CIS2, for air admittance valves | Will be equipped by the air admittance valves | B |
| 2F20 | ISO-K 100 | On the North side of CIS2, for the mass spectrometer | Will be equipped by the mass spectrometer | B |
| 2F21 | ISO-K 200 | On the underside of CIS2, for TMP (2,300 l/s) | TMP mounted for the test vacuum pumping | B |
| 2F22 | ISO-K 100 | On the underside of CIS2, for roughing pumping | Will be equipped by ISO 100 soft start valve | B |

Table 10: Positions and specifications of CIS2 vacuum flanges. Equivalent solution to ISO-K and ISO-F flange types, compatible with HV, can be provided.

# Summarised design and functional requirements for the CIS

The following tables provide a summary of the contractual requirements. The total scope of the contract comprises all the requirements stated or implied in the foregoing text, whether or not included in this summary.

## Layout and optomechanical configuration

| N° | TC N° |  | Requirement Text | Verification By |
| --- | --- | --- | --- | --- |
| Layout and optomechanical configuration |  |  |  |  |
| R1-01 | REQ-022218 | A | The coordinated final design of all components shall maintain the optical trajectory of the laser beam exactly as defined in the 3D model. NOTE: The 3D model will be provided to the Supplier after the contract signature. | R - review of design |
| R1-02 | REQ-022219 | A | The overall external dimensions of the CIS chambers shall not exceed those shown in the conceptual design as defined in the attached drawing package. | R - review of design, T - test |
| R1-03 | REQ-022220 | A | The CIS chamber bodies, doors, lids, flanges and interconnecting tubes shall be manufactured from stainless steel (304L). | R - review of design |
| R1-04 | REQ-022221 | A | All elements of the detailed design of the CIS, including optomechanics and vacuum controls, shall be mutually coordinated as these form one functional system. | R - review of design |
| R1-05 | REQ-022222 | A | The Supplier shall complete FEM analysis of each final CIS chamber design to demonstrate structural stability resulting in deformations of walls and doors less than 1 mm when the chambers are pumped down from atmospheric pressure. | A – analysis |
| R1-06 | REQ-022223 | A | Certain flanges on the CIS chambers shall have a co-axial tolerance of ±1 mm or better and angular tolerance of ±0.5 degrees or better with respect to their ideal axis; for reasons of critical geometry. NOTE: Other flanges can have standard production tolerance. The required tolerance band for each flange is specified in the Vacuum Flange Schedule (Section 10). | R - review of design, T - test |
| R1-07 | REQ-022224 | A | The CIS vacuum chambers and their internal structures shall be positioned with a tolerance of ±1 mm in each axis, with respect the local coordinate system in the L4 hall. The final design shall respect the necessity to position the chambers with this relative accuracy. | R - review of design |
| R1-08 | REQ-022225 | A | The CIS detailed design shall include full coordination with the interfaces of adjacent systems, especially primary vacuum lines, turbomolecular pumps, vacuum gauges, cooling water, compressed air, etc. and with the ELI-Beamlines building, according to information supplied by the Client. | R - review of design |
| R1-09 | REQ-022226 | A | The Supplier shall develop design of both internal (vacuum) and external (on the outer surface) cable trays on the CIS chambers. | R - review of design |
| R1-10 | REQ-022228 | A | The CIS detailed design shall make all necessary allowance for transport of the chamber to its working location. This shall include provision of designated lifting points and jacking points and positions to support the chamber on rollers for lateral movement, and any special tools or aids necessary for support chassis installation inside the chamber. | R - review of design |
| R1-11 | REQ-022229 | A | The side doors of the CIS1A and CIS2 chambers shall be side-hinged For CIS1A the door must be capable of configuring the hinges on either side. | R - review of design |
| R1-12 | REQ-022230 | A | The CIS1B chamber shall be designed by the Supplier in accordance with the conceptual design that shows a specific top access lid. The Supplier may adapt the design but it must meet the need to insert the internal optomechanical structures within the chamber from the top and to be compatible with the surrounding building structure and services. | R - review of design |
| R1-13 | REQ-022231 | A | Chambers CIS1A and CIS1B shall be connected by a square vacuum tube (internal dimensions 300 by 300 mm) equipped with a bellows at one end and an optical window at the other end. This tube shall pass under optical table number 17 of the L4 laser system where access is restricted. The design shall be coordinated with the FZU. | R - review of design |
| R1-14 | REQ-022232 | A | All electrical cables for the vacuum devices as well as compressed air and cooling water pipework shall be organized into cables trays on the CIS chamber exteriors. | R - review of design |
| R1-15 | REQ-022233 | A | Tray(s) for cooling water shall be segregated from electrical cable tray(s). | R - review of design |

## Optomechanical mounts

| N° | TC N° |  | Requirement Text | Verification by |
| --- | --- | --- | --- | --- |
| Optomechanical mounts |  |  |  |  |
| R2-01 | REQ-022234 | A | The optomechanical mounts shall be designed to support the respective optical components with dimensions as given in Table 2, Table 3, and Table 4. The nominal position and orientation of the optical components in space shall be as in the 3D model. The required extent of movements shall be as in Table 1. NOTE: The 3D model will be provided to the Supplier after the contract signature. | R - review of design, T - test |
| R2-02 | REQ-022235 | A | The optomechanical mounts inside each CIS chamber shall be attachable to the optical support chassis and shall be able to be pre-positioned (without engaging the actuators) with precision of ±1 mm. | R - review of design, T - test |
| R2-03 | REQ-022236 | A | Each individual optomechanical mount shall provide short-term and long-term angular mechanical stability (drift) of the optical element better than 1 µrad (in a thermally stabilized environment). | T - test |
| R2-04 | REQ-022237 | A | FEM analysis of the optomechanical mounts shall be carried out by the Supplier. | A - analysis |
| R2-05 | REQ-022238 | A | The mounts shall not exhibit any resonant frequency below 30 Hz, and also shall not have resonant frequency close to 50 Hz. | A - analysis |
| R2-06 | REQ-022239 | A | The principal material of the optomechanical mounts shall be aluminium so as to minimise the weight of the mount for handling purposes. | R - review of design |
| R2-07 | REQ-022240 | A | The mounts design shall feature integration of electrical actuator including wirings for electrical actuators. | R - review of design |
| R2-08 | REQ-022241 | A | The manufacturing design of the mounts and any other components inside each CIS chamber shall avoid any trapped volumes of air, e.g. the mounting holes shall not be blind tapped. | R - review of design |
| R2-09 | REQ-022242 | A | All actuators used in the optomechanics shall be certified for vacuum at least 10-6 mbar. | R - review of design |
| R2-10 | REQ-022243 | A | All edges of the optomechanical mounts shall be chamfered. | R - review of design |
| R2-11 | REQ-022244 | A | All cables for electrical actuators shall be organized into cable trays both inside the chamber and on the chamber exterior. | R - review of design |
| R2-12 | REQ-022245 | A | Cable from each electrical actuator shall be led to the exterior by no more than one electrical vacuum feedthrough, using the flanges identified in Section 10 Vacuum Flange Schedule. | R - review of design |
| R2-13 | REQ-022246 | A | Best practice shall be followed in the design, type choice and implementation of the vacuum electrical feedthroughs. | R - review of design |

## Vacuum components

| N° | TC N° |  | Requirement Text | Verification by |
| --- | --- | --- | --- | --- |
| Vacuum components |  |  |  |  |
| R3-01 | REQ-022247 | A | The CIS shall be designed for vacuum 10-7 mbar. This pressure shall be achievable with the optomechanical mounts installed in the chambers, using a turbomolecular vacuum pump with pumping speed at least 2,300 l/s permanently installed on each chamber CIS1B and CIS2, in a time comparable to 24 hours. | A – analysis, T - test |
| R3-02 | REQ-022248 | A | A pressure 10-6 mbar shall be achieved, by the turbomolecular vacuum pumps permanently installed on the CIS1B and CIS2 chambers, in less than 8 hours from start of pumping at 10-2 mbar pressure. | A – analysis, T - test |
| R3-03 | REQ-022249 | A | The CIS chambers shall provide structural stability resulting in deformations less than 1 mm of its walls and doors when the chambers are pumped down from atmospheric pressure. | A – analysis, T - test |
| R3-04 | REQ-022250 | A | All circular flanges on the CIS chambers (except for TMPs) shall be dimensioned according to the ISO-K standard, or equivalent solution compatible with HV, and will be sealed by single fluoroelastomer O-rings (see Section 10, Vacuum flange schedule). | R - review of design |
| R3-05 | REQ-022251 | A | The TMPs shall be mounted to ISO-F (or equivalent) flanges. | R - review of design |
| R3-06 | REQ-022252 | A | The large rectangular lid of CIS1B and the large door of CIS2 shall be sealed by double fluoroelastomer O-rings with the interspace evacuated. The Supplier shall develop the design of this arrangement. | R - review of design |
| R3-07 | REQ-022253 | A | All materials used inside the CIS shall comply with the regime of highly clean vacuum, free of hydrocarbons and other volatile molecules. Special care must be taken in using any plastics (e.g. insulation of electrical wires) which have to be vacuum compatible for pressures better than 10-7 mbar, e.g. Kapton® or equivalent. | R - review of design, A - analysis |
| R3-08 | REQ-022254 | A | The vacuum gauges and other elements of the CIS shall make it possible to carry out precise measurement of vacuum level down to circa 10-8 mbar. In addition atmospheric gauges, for machine safety, shall be provided. | R - review of design, I - inspection |
| R3-09 | REQ-022255 | A | The CIS2 shall be equipped with a quadrupole mass spectrometer making possible on-line measurements of cleanliness of the vacuum. | R - review of design, I - inspection |
| R3-10 | REQ-022256 | A | The CIS shall be equipped with dedicated flanges for electrical feedthroughs. | R - review of design, I - inspection |
| R3-11 | REQ-022257 | A | The Supplier shall develop detailed and rigorous procedures of cleaning the interior of the vacuum chambers and of all optomechanical components, consistent with Section 7. | R - review of design |
| R3-12 | REQ-022258 | A | The procedures of cleaning shall be reviewed and agreed by the Client before use. | R - review of design |
| R3-13 | REQ-022259 | A | The turbomolecular pumps shall be water cooled; piping of the cooling water to the pumps shall be included in the scope of detailed design and supply. The Supplier shall design, supply and install all cooling water pipework to and from the common supply point (location to be agreed but not more than 10m away) including all pipe supports, fittings, manual isolation valves, etc. as necessary for a functional system. | R - review of design, I - inspection |
| R3-14 | REQ-022260 | A | All vacuum valves shall be equipped with electro-pneumatic actuator with position indicator for both end positions. | R - review of design |
| R3-15 | REQ-022261 | A | The Supplier shall design, supply and install all compressed air distribution pipework from the common supply point to every relevant consumer, including all pipe supports, fittings, manual isolation valves, etc. as necessary for a functional system. NOTE: Air supply for vacuum valves and any other air operated equipment will be provided from a single common supply connection point near to floor level in each of L4b and L4c. | R - review of design, I - inspection |
| R3-16 | REQ-022262 | A | Backing vacuum for both CIS1B and CIS2 shall be provided from a single connection point on the backing vacuum system pipework in L4c. The design provided in the 3D model and drawing package is only indicative and the Supplier shall design and supply all the backing vacuum connection pipework from the common connection point in L4c. | R - review of design, I - inspection |
| R3-17 | REQ-022263 | A | The Supplier shall design, supply and install all vacuum pipework from the common supply point to both flanges equipped with double O-ring seal, including all pipe supports, fittings, manual isolation valves, etc. as necessary for a functional system. | R - review of design, I - inspection |
| R3-18 | REQ-022264 | A | The design, material selection and manufacturing of all components inside the vacuum envelope of the CIS shall follow best practice to avoid any trapped volumes of air. | R - review of design |
| R3-19 | REQ-022265 | A | The manufacturing processes, surface treatment and the selection of all materials used inside the vacuum envelope of the CIS shall follow best practice to avoid the incorporation of particulates, moisture and volatile organic compounds that would affect the ability to reach vacuum pump-down specification and/or potentially contaminate optics. | R - review of design, I - inspection |
| R3-20 | REQ-022266 | A | All wires and signal connections that cross the vacuum chamber boundary shall be provisioned by the Supplier with vacuum feedthroughs on the most appropriate flange (see Flange Schedule Table 8, Table 9, and Table 10). | R - review of design |
| R3-21 | REQ-022267 | A | Each logical device connection shall have an individual vacuum feedthrough connector. | R - review of design |
| R3-22 | REQ-022268 | A | All electrical vacuum feedthroughs shall be fully screened (360° EMC shielding) and all electrical wiring (both in vacuum and between the control cabinet and the chamber) shall be via electrically screened cables. | R - review of design |
| R3-23 | REQ-022269 | A | All flanges, feedthroughs, connectors, cables, pipes, terminals and devices shall by fully labelled with a secure, permanent method and uniquely identified according to a mutually agreed naming convention. | R - review of design, I - inspection |
| R3-24 | REQ-022270 | A | Under normal, non-emergency, conditions, venting (and purging) of CIS1B and CIS2 shall be via a pure, dry air supply that shall be introduced in a carefully controlled manner to avoid swirling air currents that could dislodge contamination (especially particulate). NOTE: Compressed dry air of ISO 8573 class [1,1,1] or higher is available from a single common supply connection point in L4c. | R - review of design |
| R3-25 | REQ-022271 | A | The Supplier shall design, supply and install distribution pipework of the pure, dry air from the common supply point in L4c to CIS1B and CIS2. | R - review of design, I - inspection |
| R3-26 | REQ-022272 | A | A burst-disk shall be provided on either chamber CIS1B or CIS2 to protect against overpressure during purging and venting with dry air. | R - review of design, I - inspection |
| R3-27 | REQ-022273 | A | During the initial stages of pumping of CIS1B and CIS2 chambers the evacuation of air shall be carefully controlled to avoid swirling air currents that could dislodge contamination (especially particulate) so the valves 3, 4 and 6 as shown in Figure 18 shall be soft-start valves. | R - review of design |

## Optical Support Chassis

| N° | TC N° |  | Requirement Text | Verification by |
| --- | --- | --- | --- | --- |
| Optical Support Chassis |  |  |  |  |
| R4-01 | REQ-022274 | A | The internal support chassis shall be mechanically de-coupled from the vacuum vessel by means of bellows so that the support posts (isolation legs) of the support chassis can be bolted directly to the laser building structure independently of the vacuum chamber supports. | R - review of design |
| R4-02 | REQ-022275 | A | The arrangement of the isolation legs and bellows shall conform to Section 2 (Scope of Work) and to a detailed design in the 3D model that will be provided by the Client after the contract signature, which will ensure no movement of the optical support chassis as the chamber pressure is varied. | R - review of design |
| R4-03 | REQ-022276 | A | The isolation legs shall be made from stainless steel of sufficient section and stiffness to ensure optical stability. | R - review of design |
| R4-04 | REQ-022277 | A | The optical support chassis for all three chambers shall be designed and manufactured to be free of internal stresses. For CIS1A the optical support chassis shall be made from stainless steel. For CIS1B the table of the support chassis shall be made from stress-relieved cast aluminium slab. NOTE: For the two chassis within CIS2 either form can be used. | R - review of design |
| R4-05 | REQ-022278 | A | The FEM simulations shall be made by the Supplier during the development of the manufacturing drawings to demonstrate the response of the four optical support chassis to vibration and temperature variation The analysis must show no resonances under 30Hz and close to 50Hz. | A - analysis |
| R4-06 | REQ-022279 | A | The Supplier shall design a set of self-contained mechanical handling aids for CIS2 so that both optomechanical mounts can be securely positioned outside the chamber and then installed horizontally and or vertically directly, e.g. by means of guide rails, into its respective support chassis position through the relevant chamber access door. | R - review of design |
| R4-07 | REQ-022280 | A | The handling system shall be suitable to operate in a clean room environment of Class 5 according to ČSN EN ISO 14644 standard (equivalent to ISO 14644, or other corresponding equivalent). NOTE: Regarding the referred to standard the FZU allows/permits also another equal solution to be offered). | R - review of design, T - test |
| R4-08 | REQ-022281 | A | The handling system shall not rely on the L4b or L4c hall cranes. | R - review of design |

## Control systems

| N° | TC N° |  | Requirement Text | Verification by |
| --- | --- | --- | --- | --- |
| Controls, personal and machine safety |  |  |  |  |
| R5-01 | REQ-022282 | A | Pumps, valves, actuators, encoders, sensors and other instrumentation shall be delivered with control systems, power and utilities connections as necessary for correct and safe operation of the CIS, and with full English-language documentation. | R - review of design, I - inspection |
| R5-02 | REQ-022283 | A | All control systems hardware shall be selected from an Approved Hardware List (to be finalized with the Supplier) to be compatible with the hardware used for the L4 laser controls. Use of custom-made hardware or devices not on the Approved Hardware List must be justified by the Supplier and approved by the Client on a case-by-case basis. NOTE: This list contains off-the-shelf control system devices that are compatible with the L4 system controls, have been verified by the Client and for which spares are maintained. | R - review of design |
| R5-03 | REQ-022284 | A | All pumps, valves, actuators, sensors and other instrumentation which has a computer interface such as RS-485 or Ethernet shall be provided with LabVIEW drivers and/or API and a basic demonstrative graphical user interface (GUI). All drivers shall follow the National Instruments’ Instrument Driver Guidelines, and shall be verified as compatible with the National Instruments Real-Time operating system (NI ETS 2014 – see KB 4LRA4IQ0) on the RMC-8354 real-time server computer, to ensure compatibility with the existing integration platform of the Client for the L4 laser and for other major laser equipment of ELI-Beamlines. NOTE: The control system of the L4 laser into which the CIS controls will be integrated is based on LabVIEW 2014 SP1. | R - review of design, I - inspection |
| R5-04 | REQ-022285 | A | All software forming part of the control system, including but not limited to drivers (R6-03) and FPGA firmware, shall be provided fully open-source in LabVIEW, and use only National Instruments’ distributed libraries and those by 3rd parties in wide circulation via the VI Package Manager. No DLLs, COM or .NET objects or any other mechanisms for calling closed-source external code shall be used. All software shall have full English-language documentation. | R - review of design, I - inspection |
| R5-05 | REQ-022286 | A | All critical control systems software shall run on embedded real-time computing hardware chosen from the Approved Hardware List (R5-02). NOTE: Basic GUIs and non-critical software that are not essential for correct and safe operation of the CIS may run on Windows (or equivalent) or Linux (or equivalent) PCs, as necessary. | R - review of design |

## Packaging and Transportation requirements

| N° | TC N° |  | Requirement Text | Verification by |
| --- | --- | --- | --- | --- |
| Packaging and transportation requirements |  |  |  |  |
| R6-01 | REQ-022290 | A | For the duration of its transport the CIS vacuum chambers and other vacuum envelope sections shall be hermetically sealed under dry air or nitrogen. The initial wrapping of all parts shall be in multiple layers of plastic film (as sheet or bags) of type specifically for use in contamination controlled areas. This clean conditions wrapping shall be further enclosed in robust outer packaging and transport crates as necessary for protection and handling during shipping to the ELI-Beamlines site. | I - inspection |
| R6-02 | REQ-022291 | A | The Supplier shall transport the components to the ELI-Beamlines facility and shall remain responsible (with appropriate insurance cover) up to the point of offloading at the ELI-Beamlines Facility loading ramp. Offloading of the CIS components at the building entrance will be carried out by fork lift truck (FLT) so the packages shall be palletised to suit FLT handling. | R - review of design, I - inspection |
| R6-03 | REQ-022292 | A | In order to fit in the lift the maximum dimensions of any component transport package shall not exceed 5.5 x 5.1x2.95 m3 and its weight shall not exceed 10 tons. NOTE: Transport through the Laser Building will involve the use of a goods lift to transition between floors. | R - review of design |

## Testing, inspection, documentation and warranty

| N° | TC N° |  | Requirement Text | Verification by |
| --- | --- | --- | --- | --- |
| Testing, inspection, documentation and warranty |  |  |  |  |
| R7-01 | REQ-022293 | A | The CIS chambers and associated vacuum components shall be cleaned and vacuum tested at the manufacturer’s works prior to acceptance for transport to ELI-Beamlines. NOTE: FZU will carry out its own spot check tests to confirm vacuum cleanliness of each of the three chambers using its own test method. | T - test |
| R7-02 | REQ-022294 | A | Vacuum pumping tests shall prove that the CIS chambers and associated vacuum components can be pumped down to 10-7 mbar and using the quadrupole mass spectrometer it shall be demonstrated that volatile organic compounds (VOCs) and other impurities are not present at the lower limit of detection of this spectrometer. | T - test |
| R7-03 | REQ-022295 | A | In addition to doors and access ports, all flanges and openings of the CIS chambers shall be provided with ISO-K (or equivalent solution compatible with HV) closures to enable the vessel to be vacuum tested prior its final installation at ELI-Beamlines. | R - review of design, I - inspection |
| R7-04 | REQ-022296 | A | Only new materials and equipment with manufacturer’s full warranty shall be used for the entire scope of supply. | I - inspection |
| R7-05 | REQ-022297 | A | All equipment shall be supplied with at least 24 months warranty, except turbomolecular pumps and control systems (vacuum control system and optomechanical electronic controls) where the minimum warranty period shall be at least 12 months and vacuum valves which shall be warranted for a minimum of 100,000 open/closed cycles. NOTE: The warranty for individual subsystems will be counted from the respective date of delivery to ELI-Beamlines. | I - inspection |
| R7-06 | REQ-022298 | A | The CIS shall be delivered with fully detailed installation and operational manuals, certification and technical documentation, in English language. | R - review of design, I - inspection |

## Safety Requirements

### General Safety Requirements

| N° | TC N° |  | Requirement Text | Verification by |
| --- | --- | --- | --- | --- |
| General Safety Requirements |  |  |  |  |
| R8-01 | REQ-022300 | A | The Supplier shall supply a Declaration of Conformity or any other equivalent document legally recognized and accepted in the Czech Republic for each product type if the appropriate legislation determines the Supplier's obligation to have a Declaration of Conformity (or the equivalent document) for the purposes of a Device sale in the Czech Republic to fulfil the requirements of 2001/95/EC directive or applicable Czech law. | I - inspection |
| R8-02 | REQ-022301 | A | The Declaration of Conformity shall contain at least the following norms: - ČSN EN 61010-1 ed. 2 Safety requirements for electrical equipment for measurement, control, and laboratory use (equivalent to EN 61010-1); - ČSN EN 61326-1 Electromagnetic compatibility (EMC) standard for measurement, control and laboratory use (equivalent to EN 61326-1). | R - review of design, I - inspection |

### Personal and Machine safety

| N° | TC N° |  | Requirement Text | Verification by |
| --- | --- | --- | --- | --- |
| Personal and Machine safety |  |  |  |  |
| R8-03 | REQ-022302 | A | A risk-assessment and failure mode and effects analysis (FMEA), to agreed standards, shall be conducted by the manufacture for the control systems in their scope of supply. Hazards to persons and the environment, and failures exceeding €10k damages or requiring more than 8 hours down-time shall be identified and considered. | R - review of design, A - analysis |
| R8-04 | REQ-022303 | A | Personnel Protection Systems and Machine Protection Systems shall be provided that provide sufficient control over the hazards and failures identified in R8-04 to a mutually acceptable risk level. Where possible, components used in these systems shall comply with IEC 61508 SIL3 level. | I - inspection |
| R8-05 | REQ-022304 | A | If risk assessment (FMEA) shows the need for control system mitigation for personnel protection, implementation shall comply with IEC 61508 SIL 3. | R - review of design |
| R8-06 | REQ-022305 | A | For R8-04 and R8-05, the hazard of asphyxiation during venting when using oxygen-depleted gases or gas mixtures shall be explicitly considered and the Supplier shall demonstrate that this risk is as low as reasonably practicable. | R - review of design |
| R8-07 | REQ-022306 | A | For R8-04 and R8-05, the hazard of persons becoming trapped in the chamber shall be explicitly considered and the Supplier shall demonstrate that this risk is as low as reasonably practicable. | R - review of design |
| R8-08 | REQ-022307 | A | For R8-04 and R8-05, the hazard of personnel injury when working on top of and around the sides of the chamber shall be explicitly considered and the Supplier shall demonstrate that this risk is as low as reasonably practicable (e.g., hand rails and ladder elements included in design in accordance with Czech Republic Government Resolution NV 362/2005 Annex 1). | R - review of design |
| R8-09 | REQ-022309 | A | All exposed metal parts (chamber, tubing, valves, etc.) shall be electrically bonded to the facility ground for machine and personnel safety. | R - review of design, I - inspection |

# List of Abbreviations

| Abbreviation | Meaning |
| --- | --- |
| A | Analysis (as a verification method; see details in the Annex-III) |
| API | Application Programming Interface |
| BiSS | Bidirectional / Serial / Synchronous |
| CIS | Compressor Image relay System |
| COTS | Commercial off the shelf |
| DN | Diameter Nominal |
| ELI | Extreme Light Infrastructure |
| FEM | Finite Element Method |
| FMEA | Failure Modes and Effects Analysis |
| FPGA | Field-Programmable Gate Array |
| FZU | Fyzikální ústav, Institute of Physics (IoP), Academy of Sciences of the Czech Republic |
| GUI | Graphical User Interface |
| HV | High Vacuum (approximately 10-6 - 10-8 mbar) |
| I | Inspection (as a verification method; see details in the Annex-III) |
| IoP | Institute of Physics (in Czech Fyzikální ústav, FZU) |
| ISO | International Standards Organization (here ISO-K and ISO-F vacuum flanges) |
| L x W x H | Length x Width x Height |
| L4 | Laser system No 4 of ELI-Beamlines |
| MSS | Machine Safety System |
| N/A or N.A. | Not Applicable |
| NI | National Instruments |
| OM | Optomechanics |
| PC | Personal Computer |
| PLC | Programmable Logic Controller |
| PW | Petawatt |
| NCR | Non-Conformity Report (see details in the Annex-III) |
| R | Review of Design, Review (as a verification method; see details in the Annex-III) |
| RF | Radio Frequency |
| RIO | Reconfigurable I/O |
| RMS | Reference Mechanical System |
| RSD | Requirement Specification Document |
| RTD | Resistance Temperature Detector |
| RTS | Real-time Server |
| SCADA | Supervisory Control and Data Acquisition |
| SI | International System of Units |
| SIL | Safety Integrity Level |
| T | Test (as a verification method; see details in the Annex-III) |
| TC | Teamcenter (Product Lifecycle Management System) of ELI-Beamlines |
| TMP | Turbo Molecular Pump |
| TVS | Transient Voltage Suppression |
| CDR | Critical Design Review (see details in the Annex-III) |
| VCD | Verification Control Document (see details in the Annex-III) |
| VCS | Vacuum Control System |
| VOC | Volatile Organic Compounds |
| VP | Verification Plan (see details in the Annex-III) |

# Reference drawings and diagrams

| TC ID No. | Title of document / File | Description |
| --- | --- | --- |
| 00158688/B | L4_CIS_Vacuum_Controls_Overview_V2_12-Dec-2017.pdf | L4 CIS Vacuum Controls diagram; L4 CIS Vacuum Volume State Machine; L4 CIS Cooling Water Overview; L4 CIS Pneumatics Design. |
| 00158689/B | L4_CIS_Simplified_Motion_Controls_V2_12-Dec-2017.pdf |  |
| 00159754/A | L4_CIS_Vacuum_scheme_11-Dec-2017.pdf | L4 CIS Vacuum scheme |
| 00158671/A | 00158671-A_3.4_ES_Drawing_Package-L4_CIS_TP17_044.pdf | Drawing package (see below) |

## Drawing Package

Detailed list of drawings included within “00158671-A_3.4_ES_Drawing_Package-L4_CIS_TP17_044.pdf” file:

| No. | TC ID No. | No. of sheets | Drawing Title | Description |
| --- | --- | --- | --- | --- |
| 1 | 00158671/A | 1 | L4 CIS Assembly | Isometric view |
| 2 | 00158671/A | 1 | L4 CIS Assembly | Orthographic view |
| 3 | 00158671/A | 1 | CIS1A |  |
| 4 | 00158671/A | 1 | CIS1A |  |
| 5 | 00158671/A | 1 | CIS1B |  |
| 6 | 00158671/A | 1 | CIS1B |  |
| 7 | 00158671/A | 1 | CIS2 |  |
| 8 | 00158671/A | 1 | CIS2 |  |