Nabídka ICZ180804

Testovací stanoviště motorů TITAN

Titan D 250

Technická univerzita v Liberci

Veřejná zakázka:

158_Ve_ERDF_Nový systém ovládání a řízení motorové brzdy (FS 008)

Nabídka ICZ180804

Testovací stanoviště motorů TITAN

Titan D 250

Technická univerzita v Liberci

Veřejná zakázka:

158_Ve_ERDF_Nový systém ovládání a řízení motorové brzdy (FS 008)

HORIBA GmbH, organizační složka

Technická univerzita v Liberci

Studentská 2

461 17 LIBEREC

Nabídka ICZ180804

Předkladatel nabídky			HORIBA, GmbH, organizační složka

Průmyslová 1306/7

102 00 Praha 10

Odpovědná osoba:			Ing. Václav Bubeník, vedoucí organizační složky

Nadřízený statutární orgán:	HORIBA GmbH

Kaplan-Strasse 5

A-3430 Tulln, Rakousko

Telefon:				XXXXXXXXXXXXXX

E-mail:					XXXXXXXXXXXXXXXX

IČ:					49366904

DIČ:					CZ49366904

Bankovní spojení:			Česká spořitelna, a.s.

Číslo účtu : 845572/0800 (CZK)

Zápis v Obchodním rejstříku:		Městský soud v Praze, oddíl A, vložka 7941

Datum nabídky:				8. 8. 2018

Dodací a platební podmínky

Cenové podmínky: 			DPA  Technická univerzita v Liberci

cena bez DPH  		7 779 900,- Kč

DPH 			1 633 779,- Kč

cena s DPH		9 413 679,- Kč

Dodací lhůta:				Do 6 měsíců ode dne účinnosti smlouvy

Platba:	100 % částky do 21 kalendářních dnů od doručení faktury kupujícímu. Platba se uskuteční na účet společnosti HORIBA, GmbH, organizační složka na základě vystavené faktury.

Záruční doba:	24 měsíců

Země původu:				EU

Platnost nabídky:			90 dní

Nabídku vypracoval:			Ing. Václav Bubeník

Nabídka splňuje všechny zadané podmínky výběrového řízení.

Obsah

1	Obsah	4

2	Cenová specifikace	6

3	SCOPE OF SUPPLY / SCOPE OF SERVICES	8

3.1	Dynamometer	8

3.1.1	Load unit DYNAS3 LI 250 1000070739	8

3.1.2	Custom Frame for DYNAS3 1000013603	13

3.1.3	Measuring and control cables (Set 1a) for DYNAS3 (20m) 1300023793	14

3.1.4	Measuring and control cables (Set 2a) for DYNAS3 (20m) 1000057747	14

3.1.5	Measuring cables (Set 3) for DYNAS3 with HBM-flange (20m)  1000061305	14

3.1.6	Power cable 4G70 460V for Dynas3 LI225/250/265HP, HD255,  HT250 and HS180 1000058552	15

3.1.7	Calibration unit for slotted weights for LI145/225/250 1000010997	16

3.1.8	Copper Strip for Mass Connection 1000002981	16

3.2	Dyno and Engine Controller	17

3.2.1	Digital Test Stand Controller SPARCE  1000012971	17

3.2.2	E-Gas Potentiometer Simulator 1000011123	20

3.2.3	Cable Set for E-Gas Potentiometer Simulator (20m)  1000012855	22

3.3	Measuring Equipment + Instrumentation	23

3.3.1	Voltage and Current Measuring Module SC16V 1000006843	23

3.3.2	BNC Shunt Adaptor for current measurement 1000089770	24

3.3.3	Temperature Measuring Module SDT16 1000078726	25

3.3.4	Temperature Measuring Module STh16-K 1000078727	27

3.3.5	Gauge pressure sensor ECO series 1300014752	29

3.3.6	Pressure Measuring Module SDP8 1000045153	31

3.3.7	Clean Break Coupling 1-Channel Set 1000013333	32

3.3.8	4-Channel Counter Module INC4	33

3.3.9	DSUB Connector for INC4	34

3.3.10	Wall bracket 2.5m for boom-box 1000075710	35

3.3.11	Boom Box 7U 1000012386	37

3.3.12	Cable Set Measuring Module Type D/S 0.3m 1000013434	38

3.3.13	Customized cable length for Measuring Box / Boom Box 1000013436	38

3.3.14	Analog Data Acquisition module LAM002 1000051746	39

3.3.15	Cable Set for Analogue Data Acquisition Module LAM002 1000012213	40

3.3.16	Built-in Housing for LAM002 1000007032	41

3.3.17	Weather Station with 15m cable 1000012506	43

3.3.18	Customized I/O	46

3.4	Automation System	48

3.4.1	STARS Engine Workstation 1300043538	48

3.4.2	STARS basic test stand I/O 1000004904	50

3.4.3	STARS Engine Workstation License 1000012410	52

3.4.4	STARS Device Driver License Opacimeter AVL 439 1000007055	56

3.4.5	STARS Device Driver Software License Fuel Flow Meter AVL 1000007061	58

3.4.6	STARS Device Driver AVL Indicom 1000074911	59

3.4.7	STARS Device Driver ASAM ASAP3 1000007064	61

3.4.8	Installation STARS Device Driver HORIBA MEXA ONE LAN 1000007058	63

3.4.9	Installation STARS Device Driver HORIBA MEXA 1000SPCS 1000044868	64

3.4.10	Installation STARS Device Driver HORIBA MDLT 1300T 1000007062	65

3.4.11	Installation STARS Device Driver HORIBA PTFM 1000 1000059149	66

3.4.12	Device Driver Software License Blow-By Meter 442 for STARS Test Stand Automation 1000007044	67

Technical Data	67

3.4.13	STARS RS 232 Serial  Board 1300033783	67

3.4.14	STARS High Speed CAN Interface 1300033780	67

3.5	Operators Place	68

3.5.1	Measuring and Control Cabinet Type T/D 1300042242	68

3.5.2	Central Safety Module SZM light 1000112842	71

3.5.3	Emergency Stop Button (10m) 1300006412	72

3.5.4	Emergency Stop Button (20m) 1300006417	73

3.5.5	Multiple Socket Outlet EMC CEE 7-4 (8m) 1000010550	74

3.5.6	Power Cable for Measuring and Control Cabinet T/D 1000067281	75

3.6	Services	76

3.6.1	Dokumentace papírová i na CDROM 1000010462	76

3.6.2	Instalace na místě	77

3.6.3	Uvedení do provozu	77

3.6.4	STARS  Software-uvedení do provozu na místě	77

3.6.5	Zaškolení min. 8 osob a předání	78

Cenová specifikace

| Pos. | Basic Quantity | Option Quantity | Product Number | Name | Price in EUR | Option in EUR |
| --- | --- | --- | --- | --- | --- | --- |
| 3.1 | Dynamometer |  |  |  |  |  |
| 3.1.1 | 1 |  | 1000070739 | DYNAS3 LI 250 | Incl. |  |
| 3.1.2 | 1 |  | 1000013603 | Custom Frame for DYNAS3 | Incl. |  |
| 3.1.3 | 1 |  | 1300023793 | Measuring and control cables (Set 1a) for DYNAS3 (20m) | Incl. |  |
| 3.1.4 | 1 |  | 1000057747 | Measuring and control cables (Set 2a) for DYNAS3 (20m) | Incl. |  |
| 3.1.5 | 1 |  | 1000061305 | Measuring cables (Set 3) for DYNAS3 with HBM-flange (20m) | Incl. |  |
| 3.1.6 | 20 |  | 1000058552 | Power cable 4G70 460V for Dynas3 LI225/250/265HP, HD255, HT250 and HS180 | Incl. |  |
| 3.1.7 | 1 |  | 1000010997 | Calibration unit for slotted weights for LI145/225/250 | Incl. |  |
| 3.1.8 | 40 |  | 1000002981 | Copper Strip for Mass Connection | Incl. |  |
| 3.2 | Dyno and Engine Controller |  |  |  |  |  |
| 3.2.1 | 1 |  | 1000012971 | Digital Test Stand Controller SPARCE with fixed operator panel | Incl. |  |
| 3.2.2 | 1 |  | 1000011123 | E-Gas Potentiometer Simulator | Incl. |  |
| 3.2.3 | 1 |  | 1000012855 | Cable Set for E-Gas Potentiometer Simulator (20m) | Incl. |  |
| 3.3 | Measuring Equipment + Instrumentation |  |  |  |  |  |
| 3.3.1 | 1 |  | 1000006843 | Voltage and Current Measuring Module SC16V | Incl. |  |
| 3.3.2 | 8 |  | 1000089770 | BNC Shunt Adaptor for current measurement | Incl. |  |
| 3.3.3 | 1 |  | 1000078726 | Temperature Measuring Module SDT16 | Incl. |  |
| 3.3.4 | 1 |  | 1000078727 | Temperature Measuring Module STh16-K | Incl. |  |
| 3.3.5 | 10 |  | 1300014752 | Gauge pressure sensor ECO series | Incl. |  |
| 3.3.6 | 2 |  | 1000045153 | Pressure Measuring Module SDP8 | Incl. |  |
| 3.3.7 | 10 |  | 1000013333 | Clean Break Coupling 1-Channel Set | Incl. |  |
| 3.3.8 | 1 |  |  | 4-Channel Counter Module INC4 | Incl. |  |
| 3.3.9 | 1 |  |  | DSUB Connector for INC4 | Incl. |  |
| 3.3.10 | 1 |  | 1000075710 | Wall bracket 2.5m for boom-box | Incl. |  |
| 3.3.11 | 1 |  | 1000012386 | Boom Box 7U | Incl. |  |
| 3.3.12 | 4 |  | 1000013434 | Cable Set Measuring Module Type D/S 0.3m | Incl. |  |
| 3.3.13 | 1 |  | 1000013436 | Customized cable length for Measuring Box / Boom Box | Incl. |  |
| 3.3.14 | 1 |  | 1000051746 | Analog Data Acquisition module LAM002 | Incl. |  |
| 3.3.15 | 1 |  | 1000012213 | Cable Set for Analogue Data Acquisition Module LAM002 | Incl. |  |
| 3.3.16 | 1 |  | 1000007032 | Built-in Housing for LAM002 | Incl. |  |
| 3.3.17 | 1 |  | 1000012506 | Weather Station with 15m cable | Incl. |  |
| 3.3.18 | 1 |  |  | I/O Box | Incl. |  |
| 3.4 | Automation System |  |  |  |  |  |
| 3.4.1 | 1 |  | 1300043538 | STARS Engine Workstation | Incl. |  |
| 3.4.2 | 1 |  | 1000004904 | STARS basic test stand I/O | Incl. |  |
| 3.4.3 | 1 |  | 1000012410 | STARS Engine Workstation License | Incl. |  |
| 3.4.4 | 1 |  | 1000007055 | STARS Device Driver License Opacimeter AVL 439 | Incl. |  |
| 3.4.5 | 1 |  | 1000007061 | STARS Device Driver Software License Fuel Flow Meter AVL | Incl. |  |
| 3.4.6 | 1 |  | 1000074911 | STARS Device Driver AVL Indicom | Incl. |  |
| 3.4.7 | 1 |  | 1000007064 | STARS Device Driver ASAM ASAP3 | Incl. |  |
| 3.4.8 | 1 |  | 1000007058 | Installation STARS Device Driver HORIBA MEXA ONE LAN | Incl. |  |
| 3.4.9 | 1 |  | 1000044868 | Installation STARS Device Driver HORIBA MEXA 1000SPCS | Incl. |  |
| 3.4.10 | 1 |  | 1000007062 | Installation STARS Device Driver HORIBA MDLT 1300T | Incl. |  |
| 3.4.11 | 1 |  | 1000059149 | Installation STARS Device Driver HORIBA PTFM 1000 | Incl. |  |
| 3.4.12 | 1 |  | 1000007044 | STARS Device Driver License BlowBy Meter AVL 442 | Incl. |  |
| 3.4.13 | 1 |  | 1300033783 | STARS RS 232 Serial Board | Incl. |  |
| 3.4.14 | 1 |  |  | STARS High Speed CAN Interface | Incl. |  |
| 3.5 | Operators Place |  |  |  |  |  |
| 3.5.1 | 1 |  |  | Measuring and Control Cabinet Type T/D | Incl. |  |
| 3.5.2 | 1 |  | 1000112842 | Central Safety Module SZM light | Incl. |  |
| 3.5.3 | 1 |  | 1300006412 | Emergency Stop Button (10m) | Incl. |  |
| 3.5.4 | 2 |  | 1300006417 | Emergency Stop Button (20m) | Incl. |  |
| 3.5.5 | 1 |  | 1000010550 | Multiple Socket Outlet EMC CEE 7-4 (8m) | Incl. |  |
| 3.5.6 | 1 |  | 1000067281 | Power Cable for Measuring and Control Cabinet T/D | Incl. |  |
| 3.6 | Services |  |  |  |  |  |
| 3.6.1 | 1 |  | 1000010462 | Dokumentace papírová i na CDROM | Incl. |  |
| 3.6.2 | 1 |  |  | Zabalení, doprava a instalace | Incl. |  |
| 3.6.3 | 1 |  |  | Uvedení do provozu | Incl. |  |
| 3.6.4 | 1 |  |  | Uvedení do provozu softwaru STARS | Incl. |  |
| 3.6.5 | 1 |  |  | Zaškolení min. 8 osob a předání | Incl. |  |
|  |  |  |  | Celkem 7 779 900,- bez DPH | CZK |  |

SCOPE OF SUPPLY / SCOPE OF SERVICES

Dynamometer

Load unit DYNAS3 LI 250	1000070739

This DYNAS3 load unit consists of the AC machine and a variable frequency drive. The entire system is designed and tested to interact with any kind of engine development, performance or function test stand. The application spectrum ranges from simple steady-state test stands to sophisticated high dynamic test stands.

Figure 1: DYNAS3-Asynchronous machine with measuring flange T40B
(design similar to illustration)

Benefits

- Superior dynamic characteristics such as:

- High speed torque rise time

- Low moment of inertia

- High overload capacity for high acceleration rates

- Wide range of usage such as:

- Large connectable overhung masses

- Wide speed range

- High reliability due to:

- Low wear and self-isolating hybrid bearings

- Lifetime lubricated bearings

- High-precision torque measurement because:

- Torque measured directly at the shaft connection for dynamically correct measurement

- Torque measuring flange with integrated evaluation unit

- High mechanical overload capacity

- Low-noise feed back to electrical power system through:

- Practically sinusoidal line currents

- Integrated power filter

- Compact dimensions for a smaller test cell size due to:

- Compact AC induction motor

- Compact frequency drive cabinet

Figure 2: DYNAS3 LI 250 with measuring flange T40B,
view without adapter flange, without protection cover

Figure 3: Measuring flange T40B, connection adapter flange

Figure 4: Adapter flange T40B

Figure 5: VFD cabinet

| Scope of Supply | Quantity |
| --- | --- |
| Asynchronous machine DYNAS3 - with terminal box, cable glands - with cord grip | 1 |
| Torque measuring flange T40B (T12 optional) with fixture, Adapter flange and Protection cover | 1 |
| Speed measuring unit | 1 |
| Frequency converter consisting of: - Supply field - Power filter - 4%-impedance - Mains and machine converter - AC-machine connection - Power Supply for measuring cabinet | 1 |
| Cabinet base 200 mm | 1 |
| Set of installation material for cabinet | 1 |
| Documentation: - Manual - Circuit diagram - Factory calibration certificate | 1x CD, 1x paper |

| Technical Data |  |
| --- | --- |
| Directions of rotation | 2 |
| Rated torque (absorbing /driving) [Nm] | 480/442 |
| Rated power (absorbing /driving) [kW] | 250/225 |
| Rated speed (absorbing /driving) [rpm] | 4980/4860 |
| Max. speed [rpm] | 10000 |
| Overload factor (max. 60s every 10min) | 1.2 |
| Max. speed gradient up to rated speed incl. overload [rpm/s] | 15600 |
| Moment of inertia [kgm2] | 0.36 |
| Max. deviation, speed [rpm] | ±1 for frequencies < 10 Hz |
| Overhung mass at nmax at lever arm 60mm from the torque flange [kg] | with measuring flange T40B: 10.4 optional with meas. flange T12: 9.3 |
| Dimension of AC motor: Overall weight [kg] | refer to figure 2 740 |
| Dimension of cabinet Weight [kg] | refer to figure 5 950 |
| Approval machine / drive cabinet | CE / CE |
| Refer to technical specification | DYNAS3 for Engine Test Stands |

| Technical Data | Measuring Flange T40B |
| --- | --- |
| Measuring range [Nm] | 1000 |
| Accuracy class | 0.05 |
| Linearity error including hysteresis, related to the nominal (rated) sensitivity [%] | < 0.03 |
| Effect of temperature per 10 K in the nominal (rated) temperature range +10 .... +70°C on the output signal, related to the actual value of the signal spread [%] on the zero signal, related to the nominal (rated) sensitivity [%] | < 0.05 < 0.05 |
| Refer to technical specification | Data sheet T40B by HBM |

Figure 6: power diagram in absorbing operation

| Interface |  |
| --- | --- |
| Power supply converter cabinet: - - Connected Power [kVA] - - Short-circuit power [MVA] - - Line voltage [V] AC - - Voltage tolerance [%] - - Frequency [Hz] | 289 ≥ 11.6 3/PEN 400 +10/-15 50 |
| Converter cabinet: - Cooling air requirement [m³/h] - Cooling air temperature [°C] - Max. relative air humidity (non-condensing) - Power loss [kW ] - Conditions for the mains connection | 4000 +5…+40 85% at 28 °C ≤ 15 Conditions for the mains connection for converter cabinets *) |
| Dynamometer: - - Cooling air requirement [m³/h] - - Cooling air temperature [°C] (non-condensing) | 900 +5…+40 |
| Interface to test stand controller | CAN-Bus |
| Power supply outlet for measuring system | 1/N/PE 230V AC, 16A |
| interface for fixing the asynchronous machine | refer to figure 2 4 x M16 |
| Shaft height | refer to figure 2 |
| Threaded holes for vibration sensors - at A-site M 8 x 10 mm - at B-site M 8 x 10 mm | 2 2 |
| Shaft connection | Figure 3 or 4 selectively |
| Cable entry mains and motor cables | from the bottom |

Note:

- Depending on shaft design the adapter flange is necessary or not.

- Feeding from the top is possible as an option (additional cabinet necessary)

- *) To ensure the operating reliability of the existing mains where the drive should be connected, the requirements of the HORIBA specification "Conditions for the mains connection for converter cabinets" have to be observed.
Otherwise it is not possible to ensure:

- Compliance with rules and regulations of the responsible electric power company

- Avoidance of mains interference, mains resonance and system failures.

- With unknown connection conditions or doubt we recommend to perform a power network analysis. This can be done on request by HORIBA as an additional service.

Custom Frame for DYNAS3	1000013603

heightening of a HORIBA dynamometer DYNAS3 LI145, LI225/250, LI350 or HT250 to a customized center height

Figure 1: frame for HORIBA dynamometer DYNAS3

| Scope of Supply | Quantity |
| --- | --- |
| Design of a frame in 3D CAD including frequency analysis related to the speed range of the dynamometer | 1 |
| Frame for dynamometer | 1 |
| Set of mounting bolts for fixing the frame and the dynamometer | 1 |
| Documentation: - assembly drawing | 1x CD, 1x paper |

| Technical Data |  |
| --- | --- |
| Possible shaft height min. / max. [mm] | 500 - 900 |
| Possible T-Slots distance, e.g. for mounting plate by customer min. / max. [mm] | 500 - 1100 |
| Color | RAL 7035 light grey |
| Material | steel |

| Interface |  |
| --- | --- |
| T Slot | DIN 650 22 H12 |
| Customer’s data - HORIBA DYNAS3 type - T-Slots distance [mm] - Shaft height [mm] |  |

Note:

- If the intermediate frame is purchased together with the dynamometer, the frame will be mounted.

Measuring and control cables (Set 1a) for DYNAS3 (20m)	1300023793

| Scope of Supply | Quantity |
| --- | --- |
| Measuring and control cable set between AC motor and VFD cabinet | 1 |

| Technical data |  |
| --- | --- |
| Cable length [m] | 20 |

Measuring and control cables (Set 2a) for DYNAS3 (20m)	1000057747

| Scope of Supply | Quantity |
| --- | --- |
| Measuring and control cable set between SPARC and VFD cabinet with CAN-interface | 1 |

| Technical data |  |
| --- | --- |
| Cable length [m] | 20 |

Measuring cables (Set 3) for DYNAS3 with HBM-flange (20m) 	1000061305

| Scope of Supply | Quantity |
| --- | --- |
| Torque/Speed actual value cables between ASM and meas. cabinet | 1 |

| Technical data |  |
| --- | --- |
| Cable length [m] | 20 |

Power cable 4G70 460V for Dynas3 LI225/250/265HP, HD255, 
HT250 and HS180	1000058552

Power cable to pass 3-fold parallel between frequency converter and AC machine.

Delivery takes place as material sold by the meter at a cable drum.

| Scope of Supply | Quantity |
| --- | --- |
| Power cable, drum (material sold by the meter) | 1 |

| Technical Data |  |
| --- | --- |
| Overall length on drum [m] | length acc. to price sheet x 3 |
| Effective cable length (usable free length) between frequency converter and AC machine [m] | length acc. to price sheet – 3m (approx. 3m will be needed for the electrical connection at the AC machine and inside the converter cabinet) |
| Approval | CE |
| Cable type | power cable, shielded wire color code VDE 0293 |
| Nominal voltage Cross section [mm²] Bending radius fix installed [mm] Bending radius flexible installed [mm] External diameter [mm] | 460 4 x 70 260 868 44 |

| Interface |  |
| --- | --- |
| Number of parallel to be laid cables | 3 |

Note:

- The power cable is especially dimensioned for the operation of this Dynas3 machine. Only the approved cable type may be used.

- The cables have to be cut into sections, to be passed and to be lagged with cable shoes (part of the electrical installation)

Calibration unit for slotted weights for LI145/225/250	1000010997

for calibration of the torque measuring chain (unit Nm) with radial stress free torque introduction for AC Motor LI145 / LI225 / LI250

Figure 1: calibration unit for 2 direction of rotation

| Scope of Supply | Quantity |
| --- | --- |
| Testing lever | 2 |
| Weight tray for weight washer | 2 |
| Documentation: - Test certificate - Drawing Z008941002 | 1x CD, 1x paper |

| Technical Data |  |
| --- | --- |
| Suitable only for shaft height [mm] | > 300 |

Note:

- Without weights

- The Dynas3 load unit are equipped with a blocking system.

Copper Strip for Mass Connection	1000002981

For potential equalization between inverter and electrical motor/frame (if not provided by local installation conditions)

| Scope of Supply | Quantity |
| --- | --- |
| Single blank copper strip | 1 |

| Technical Data |  |
| --- | --- |
| Dimensions [mm] | 15 x 1 |
| Overall length [m] | length acc. to price sheet |
| Effective cable length (usable free length) between frequency converter and AC machine [m] | length acc. to price sheet – 3m (approx. 3m will be needed for the electrical connection at the AC machine and inside the converter cabinet) |

Dyno and Engine Controller

Digital Test Stand Controller SPARCE 	1000012971

Controls DT, WT, Dynas3 dynamometer  or tandem applications of the engine test stand. For control of torque, speed, throttle position, or a free selectable parameter x.

Monitors speed, torque, power and additional 3 limit value contacts 
(oil temperature, oil pressure, and cooling water temperature) as well as a second speed input  for detection of a mechanical shaft failure by comparing speed in front and behind the shaft connection with selectable error responses.

Manual operator panel for demand value presetting of the specimen and loading device with  selectable engineering units as well as operating language English, French, Spain or German.

Display of demand values, feedback values, status signals, error messages.

Figure 1: picture of SPARCE with control panel

| Scope of Supply | Quantity |
| --- | --- |
| 19” 4HU controller rack with power supply unit | 1 |
| PC 104 CPU module | 1 |
| PC 104 power module | 1 |
| PC 104 CAN communication module | 1 |
| PC 104 dyno and engine I/O module | 1 |
| Dyno interface board, preconfigured I/O: - 6 digital inputs - 4 digital outputs - 1 safety digital output - 1 analogue input - 3 analogue outputs - 2 pulse inputs - 1 bridge input | 1 |
| Engine interface board, preconfigured I/O: - 6 digital inputs - 4 digital outputs - 1 safety digital output - 4 analogue inputs - 2 analogue outputs - 1 pulse input | 1 |
| CAN bus interfaces, intended for - test stand automation - throttle sensor simulation - operator panel and additional I/O functions - frequency converter Dynas3 - open simulation - not used | 6 |
| RS232 service interface | 1 |
| SPARC operating panel - 7” LCD for numerical display of torque, speed, throttle position, power and control mode - 3 digital demand value potentiometers - 5 pre-configured function keys - 8 available function keys (F1-F8), programmable - emergency-stop button | 1 |
| RS232 service interface | 1 |
| Line cord | 1 |
| Basis controller Software (M, n, …) | 1 |
| Flash card | 1 |
| Documentation: - Operating Manual - Commissioning Manual (including interface diagrams) - System Manual | 1x CD, 1x paper |

| Technical Data |  |
| --- | --- |
| Color | RAL 7035 light grey |
| Material front panel | aluminum (AlMg) |
| Supply voltage [V AC] | 110 / 230 |
| Supply voltage deviation [V AC] | 100 - 240 |
| Mains frequency [Hz] | 50 / 60 |
| Max. mains frequency deviation [%] | ± 5 |
| Power consumption nominal [W] | 150 |
| Connector type | IEC 320 C14 (male) |
| Steady-state control deviation torque [%] | with Dynas3: < ± 0.17 in relation to maximum measuring value with hydraulic brake DT-series / with eddy current dynamometer E-/ W-series: ± 1 in relation to maximum measuring value with eddy current dynamometer WT-series:  0,6 , in relation to maximum measuring value |
| Steady-state control deviation speed [rpm] | with Dynas3: ± 1 max. 0.25 °/oo in relation to maximum measuring value nmax, with hydraulic brake DT-series / with eddy current dynamometer E-/W-series: ± 10 with eddy current dynamometer WT-series: ± 3 |
| Dimension measuring rack | 4 HU / 19” acc. to DIN 41494-1 dimensions for 19” racks |
| Installation depth [mm] | 360 |
| Weight [kg] | 10 |
| Ambient conditions |  |
| Temperature range at operation [°C] | 0..40 |
| Relative air humidity | Class F (DIN 40050) max. 95% at 30 days/year, no condensing |
| Electrical connections | details see specification |
| Protection class housing acc. to EN 60529 | IP 10 (with equipped back panel) |
| Refer to technical specification | Test Stand Controller SPARCE |

| Interface |  |
| --- | --- |
| Connection to test stand automation | CAN bus, RS232 or Ethernet interface |
| AI, Analogue In | Voltage: +/- 10V differential Resolution: 16 bit Accuracy: 0.1% MBE |
| AO, Analogue Out | Voltage: +/- 10V Current: +/- 20mA @ 500 Ohm Resolution: 16 bit Quiescent state: 0 V, 0 mA |
| DI, Digital In | Galvanic isolated: 10-30 V; 0-0.8 V |
| DO Digital Out | Relay contact: 250 VAC, 8 A Quiescent state: Open |
| SO, Safety Output | Redundant relay contact: 250 VAC, 8 A Quiescent state: Open |
| Dyno speed acquisition | 1 or 2-phase encoder Max. frequency: 500 kHz Resolution: 16 bit Inputs: RS422, 5-24 V; open collector |
| Dyno torque acquisition | Max. frequency: 500 kHz Resolution: 16 bit Inputs: RS422, 5-24 V |
| Bridge input signal | DMS full bridge, 2mV/V Excitation: 10V @ 5kHz Resolution: 16 bit Accuracy: 100 ppm / C° MBE |
| Electrical connections | Details see documentation |

Note:

- If the automation system STARS is purchased, the test stand can also be operated manually by using the PC.

E-Gas Potentiometer Simulator	1000011123

for electronic simulation of the pedal sensor potentiometer of the combustion engine by the SPARCE, SPARCPowertrain or x-act digital test stand controller. The analogue and digital interface modules are completely enclosed.

Typically for installation close to the specimen on a mounting support.

Figure 1: E-Gas potentiometer simulator hardware

Figure 2: E-Gas potentiometer simulator hardware dimensions

| Scope of Supply | Quantity |
| --- | --- |
| SPARCE software package E-gas simulator | 1 |
| CAN module E-gas 2DO/4AI/4AO 2 AO for pre-setting the throttle position 2 AI for connecting the reference voltage from the ECU 2 Isolated DO for displaying idle and full throttle position | 1 |
| Default setup data record | 1 |
| E-Gas-module housing 1 Screw cover 3 EMC cable glands 9 Hole plugs (brass) 1 top hat rail (DIN) 2 mounting bracket (welded on) | 1 |
| Documentation: included in SPARC or x-act manual | 1x CD, 1x paper |

| Technical Data |  |
| --- | --- |
| Power supply BUS (electronics) Supply voltage [V DC] Max. supply volt. deviation [%] Max. current consumption [mA] | 24 -25..+30 500 |
| Power supply field Supply voltage [V DC] Max. supply volt. deviation [%] Max. current feeding [A] | 24 -25..+30 10 |
| Dimensions Housing | Refer to figure 2 |
| Dimension of mounting bracket | Refer to figure 2 |
| Dimension of slotted hole [mm] | 15 x 7 |
| Protection class housing acc. to EN 60529 | IP 66 |
| Color | RAL 7035 light grey, structure lacquer |
| Material | sheet steel |
| Approval housing, CAN-module | CE, UL/CSA |

| Interface |  |
| --- | --- |
| DO, potential-free (relais), NOC, AC 230 V, DC 30V, max. 2A | 2 |
| Analog outputs (0 - 10V DC) | 2 |
| Analog inputs (0 - 10V DC) | 2 |
| Platform | SPARCE hardware |
| Connection type | spring clamps |
| Wire cross section [mm² (AWG)] | 0.08..2.5 (28..14) |
| Bus type | CANopen |
| EMC gland size for cable diameter [mm] | M16 x 1,5 5..9 |

Note:

The later supplementation of an older controller x-act with an E-Gas Potentiometer Simulator can only be effected together with a separate order for approx. 4 to 6 hours for commissioning and upgrading of the controller.

For simulation by x-act the „old“ Wago terminal 750-307 is necessary.

Cable Set for E-Gas Potentiometer Simulator (20m)		1000012855

| Scope of Supply | Quantity |
| --- | --- |
| Power supply cable DC 24V to DC power supply unit and CANopen bus connecting cable to the test stand controller SPARC | 1 |
| ECU control and analogue cable to the connection at the specimen | 1 |

| Technical Data |  |
| --- | --- |
| Cable length power and bus cable [m] | 20 |
| Cable length ECU connection cables [m] | 5 |

Measuring Equipment + Instrumentation

Voltage and Current Measuring Module SC16V	1000006843

19” plug-in module for measuring voltage and current signals provided with CAN connection to test stand automation. Typical for installation in a measuring- or boom box

Figure 1: measuring module SC16V

| Scope of Supply | Quantity |
| --- | --- |
| Voltage and current measuring rack SC16V with 16 BNC connectors | 1 |
| Factory calibration certificate for the measuring module | 1 |
| Documentation operating manual measuring modules type S | 1x CD, 1x paper |

| Technical Data |  |
| --- | --- |
| Color | RAL 7035 light grey |
| Material front panel | aluminum (AlMg) |
| Resolution | 16 bit |
| Operating voltage [V] | +24 |
| Operating voltage deviation [V] | -9..+32 |
| Sampling rate max. per channel [Hz] | 500 |
| Voltage input range [V] | ±0.1, ±0.2, ±0.5, ±1, ±2, ±5, ±10 |
| Current input range [mA] with shunt connector 50 | ±10, ±20 |
| Input type | DC / differential |
| Dimension measuring rack | 1 HU / 19” acc. to DIN 41494-1 dimensions for 19” racks |
| Installation depth [mm] | 285 |
| Power consumption nominal [W] | 2.6 |

| Interface |  |
| --- | --- |
| Bus type | CAN bus (up to 1 Mbit/s, design to ISO 11898) |
| Electrical Connection | CAN-Bus connection and 24V DC supply each with 2x D-Sub 9p connectors |
| Connector type -X1..-X16 | BNC (RG58 or 223 / 50) |

Note:

- Measurement of current requires a 50-Ohm shunt resistor per channel.

BNC Shunt Adaptor for current measurement	1000089770

High-precision shunt resistor as a BNC gender changer for the measurement of current signals.

Able to be directly connected to the SC16V HORIBA measuring modules.

Figure 1: BNC Shunt Adaptor

Benefits:

- Confusion-proof – immediately visible from looking at the module whether a channel is used for current or voltage measurement

- Simply plug in to convert voltage measurement to current measurement – no internal changes are needed in the measurement module.

| Scope of Supply | Quantity |
| --- | --- |
| High-precision BNC shunt adapter | 1 |

| Technical Data |  |
| --- | --- |
| Resistance [] | 100 |
| Accuracy | 0.01 TK5 |
| Capacity [W] | 0.25 |
| Input Signal [mA] | ±20 |
| Output Signal [V] | ±2 |

| Interface |  |
| --- | --- |
| Connector | BNC f/m |

Temperature Measuring Module SDT16	1000078726

19” plug-in module for measuring temperature with 8 Pt100 RTDs and 8 type K thermocouples with CAN connection to test stand automation

Typical for installation in a measuring- or boom-box in combination with corresponding pluggable sensors

Figure 1: measuring rack SDT16

| Scope of Supply | Quantity |
| --- | --- |
| Temperature measuring rack SDT16 | 1 |
| 4-pin Pt100 connectors | 8 |
| 2-pin female thermo connectors | 8 |
| Factory calibration certificate for the measuring module | 1 |
| Test certificate with the result of the functional test of all measuring channels | 1 |
| Documentation operating manual measuring modules type S | 1x CD, 1x paper |

| Technical Data |  |
| --- | --- |
| Color | RAL 7035 light grey |
| Material front panel | aluminum (AlMg) |
| Resolution | 16 bit |
| Operating voltage [V] | +24 |
| Operating voltage deviation [V] | -9..+32 |
| Measuring rate recommended maximum for optimized noise reduction per channel [Hz] | 1 |
| Measuring rate, adjustable up to [Hz] | 100 |
| Measuring error, static, Pt100 -200...+850°C factory calibration on-site calibration | <1K <1.5K |
| Measuring error, static, thermocouple -200...1200°C factory calibration on-site calibration | <1K <2.2K |
| Operating temperature range [°C] | -40...65 |
| Default Measuring ranges | set to –200..+1200°C for thermocouples and 0..200°C for resistance thermometer (can be changed by customer). The accuracy applies to FS. |
| Input channel protective circuit [kV] | 3.5 |
| Dimension measuring rack | 1 HU / 19” acc. to DIN 41494-1 dimensions for 19” racks |
| Installation depth [mm] | 285 |
| Power consumption nominal [W] | 2.6 |

| Interface |  |
| --- | --- |
| Bus type | CAN bus (up to 1 Mbit/s, design to ISO 11898) |
| Electrical Connection | CAN-Bus connection and 24V DC supply each with 2x D-Sub 9p connectors |
| Pt100 connector type | 4p Lemo FFA.1E series |
| Thermocouple connector, ANSI color code | 2p miniature thermo connector type K, yellow |

Temperature Measuring Module STh16-K	1000078727

19” plug-in module for measuring temperature with 16 type K thermocouples with CAN connection to test stand automation

Typical for installation in a measuring- or boom-box in combination with corresponding pluggable sensors

Figure 1: measuring module STh16

| Scope of Supply | Quantity |
| --- | --- |
| Temperature measuring rack STh16 | 1 |
| 2-pin female thermo connectors | 16 |
| Factory calibration certificate for the measuring module | 1 |
| Test certificate with the result of the functional test of all measuring channels | 1 |
| Documentation operating manual measuring modules type S | 1x CD, 1x paper |

| Technical Data |  |
| --- | --- |
| Color | RAL 7035 light grey |
| Material front panel | aluminum (AlMg) |
| Resolution | 16 bit |
| Operating voltage [V] | +24 |
| Operating voltage deviation [V] | -9..+32 |
| Measuring rate recommended maximum for optimized noise reduction per channel [Hz] | 1 |
| Measuring rate, adjustable up to [Hz] | 100 |
| Measuring error, static, thermocouple -200...1200°C factory calibration on-site calibration | < 1K < 2.2K |
| Operating temperature range [°C] | -40...65 |
| Default Measuring ranges | set to –200..+1200°C (can be changed by customer). The accuracy applies to FS. |
| Input channel protective circuit [kV] | 3.5 |
| Dimension measuring rack | 1 HU / 19” acc. to DIN 41494-1 dimensions for 19” racks |
| Installation depth [mm] | 285 |
| Power consumption nominal [W] | 2.6 |

| Interface |  |
| --- | --- |
| Bus type | CAN bus (up to 1 Mbit/s, design to ISO 11898) |
| Electrical Connection | CAN-Bus connection and 24V DC supply each with 2x D-Sub 9p connectors |
| Thermocouple connector, ANSI color code | 2p miniature thermo connector type K, yellow |

Gauge pressure sensor ECO series	1300014752

For measurement of the relative pressure of gaseous or fluid media in the field of engine and powertrain testing in the automotive industry.

Compatible for installation in a HORIBA measuring rack type DP, SDP8 or EP8 or separately close to the specimen.

| Figure 1:Pressure sensor | Figure 2: Dimensions |
| --- | --- |

| Typical application pressure sensors |  |  |  |  |
| --- | --- | --- | --- | --- |
|  | engine | gearbox | meas. range [kPa] | mode |
| barometric pressure |  |  | 85 .. 120 | abs. |
| oil gallery |  |  | 0 ... +1000 | rel. |
| gearbox oil / automatic |  |  | 0 ... +1600 | rel. |
| intake air |  |  | -100 ... +200 | rel. |
| intake air / turbo |  |  | -100 ... +300 | rel. |
| fuel |  |  | 0 ... +600 | rel. |
| engine coolant |  |  | 0 ... +250 | rel. |
| charge-air |  |  | 0 ... +250 | rel. |
| turbo / compressor |  |  | 0 ... +250 | rel. |
| intercooler |  |  | 0 ... +250 | rel. |
| exhaust back pressure |  |  | 0 ... +100 | rel. |

Figure 3: Overview typical applications and measuring ranges

| Scope of Supply | Quantity |
| --- | --- |
| Pressure transducer ECO | 1 |
| Connector | 1 |
| Documentation datasheet with assembly instructions | 1 |

| Technical Data |  |  |  |  |
| --- | --- | --- | --- | --- |
| Measuring range [kPa] / quantity | Your choice: |  |  |  |
| No. 14651 / 14652 No. 14653 / 14654 No. 14655 / 14656 No. 14657 / 14658 No. 14659 / 14660 No. 14661 / 14662 No. 14682 / 14683 No. 14684 / 14685 No. 14686 / 14687 No. 14688 / 14689 No. 14700 / 14701 > 1000kPa types No. 14650 / 14608 customized measuring range(s): No. 14723 | 0...+10 0...+50 0...+150 0..+250 0..+400 0..+600 -10..+10 -20..+20 -50..+50 -100..+200 -100..+400 0..+1600 [xxx]..+ [xxx] | [ ] [ ] [ ] [ ] [ ] [ ] [ ] [ ] [ ] [ ] [ ] [ ] [ ] | 0..+20 0..+100 0..+200 0..+300 0..+500 0..+1000 -15..+15 -30..+30 -100..+100 -100..+300 -100..+500 0..+2000 | [ ] [ ] [ ] [ ] [ ] [ ] [ ] [ ] [ ] [ ] [ ] [ ] |
| Accuracy accuracy according to IEC 60770 limit point adjustment (non linearity, hysteresis, repeatability) | < 40 kPa: ≤ 0,25% FSO BFSL ≤ ± 0,5 % FSO ≥ 40 kPa: ≤ 0,175% FSO BFSL ≤ ± 0,35 % FSO |  |  |  |
| Long term stability at reference conditions | ≤ ± 0,1 % FSO / year |  |  |  |
| Thermal effects (Offset and Span) Nominal pressure PN [bar] Tolerance band [% FSO] in compensated range [°C] | -1 … 0 / < 0,40 / ≥ 0,40 ≤ ± 0,75 / ≤ ± 1 / ≤ ± 0,75 -20 ... 85 / 0 ... 70 / -20 … 85 |  |  |  |
| Permissible temperatures medium electronics / environment storage | -40 ... 125 °C -40 ... 85 °C -40 ... 100 °C |  |  |  |
| Measuring signal [mA] | 4...20 |  |  |  |

| Interface |  |
| --- | --- |
| Connector type | acc. to DIN EN 175301-803 |
| Pressure connection | ¼“ inside thread |

Note:

If ordered together with a measuring module type SDP8 or EP8, the transducer will be mounted into the rack by HORIBA.

Customized measuring ranges smaller +/-10kPa are not possible.

The transducers are intended for stationary pressure measurement, transmitters for measuring of dynamic pressures on request.

Pressure Measuring Module SDP8	1000045153

19” plug-in module SDP8 is a plug-in module for measuring pressures of gaseous and liquid media with CAN bus connection to test stand computer.

Typical for installation of pressure transducers type RS / RH and clean brake couplings e.g. in a measuring- or boom-box

Figure 1: Measuring Module SDP8
(with installed optional pressure transducers and optional clean brake couplings

| Scope of Supply | Quantity |
| --- | --- |
| Pressure measuring module SDP8 | 1 |
| installation place for clean brake coupling and transducers | 8 |
| Factory calibration certificate for the measuring module | 1 |
| Test certificate with the result of the functional test of all integrated measuring channels | 1 |
| Documentation operating manual measuring modules type S | 1x CD, 1x paper |

| Technical Data |  |
| --- | --- |
| Color | RAL 7035 light grey |
| Material front panel | aluminum (AlMg) |
| Resolution | 16 bit |
| Operating voltage [V] | +24 |
| Operating voltage deviation [V] | -9..+32 |
| Sampling rate max. per channel [Hz] | 500 |
| Dimension measuring rack | 1 HU / 19” acc. to DIN 41494-1 dimensions for 19” racks |
| Installation depth [mm] | 285 |
| Power consumption nominal [W] | 2.6 |

| Interface |  |
| --- | --- |
| Bus type | CAN bus (up to 1 Mbit/s, design to ISO 11898) |
| Electrical Connection | CAN-Bus connection and 24V DC supply each with 2x D-Sub 9p connectors |

Clean Break Coupling 1-Channel Set	1000013333

for installation in the pressure measuring rack SP8 in combination with pressure transducers type RL / RS / RH. For low leakage hose connection of gaseous and fluid media

Figure 1: clean Break coupling

Figure 2: dimensions of clean Break coupling

| Scope of Supply | Quantity |
| --- | --- |
| Female quick connector for panel installation | 1 |
| Lock nut | 1 |
| Seal ring G¼“ | 1 |
| Male quick connector (nipple) for hose connection | 1 |

| Technical Data |  |
| --- | --- |
| Material | steel, zinc coated |
| Sealing | NBR (Nitrile Butal Rubber) |
| Material seal ring | PVC |

| Interface |  |
| --- | --- |
| For installation in combination with pressure transducers recommended sealant | liquid acylate sealing; e.g. LOCTITE 572 (M0043499.01) |
| Connection female connector | G¼“ external thread |
| Hose connection male connector; diameter [mm] | 6 |
| Max. pressure [kPA] | 10.000 |

Note:

Matching plug-in hose for e.g. mineral based hydraulic and lubricating oils and water-oil emulsions optional available.

4-Channel Counter Module INC4

Module for acquisition of frequency, rotation, position and angle counter signals

| Parameter | Value | Remarks |
| --- | --- | --- |
| Inputs | 4 + 1 ( 9 tracks ) | 4 channels with 2 tracks (X, Y) each 1 index channel all fully conditioned inputs isolated from CAN-Bus and supply, but not mutually |
| Measurement modes | displacement, angle, events, time, frequency; velocity, RPMs |  |
| Sampling rate (CAN output) | 1 kHz / channel (max.) |  |
| Time resolution of measurement frequency stability deterioration | 33 ns <100 ppm < 5 ppm / year | counter frequency 32 MHz (primary sampling rate) |
| Resolution of data | 16 bits |  |
| Input configuration | differential |  |
| Input impedance | 100 kOhm |  |
| Switching threshold | -8 V to +10 V |  |
| Hysteresis | 0.3 V to 4 V |  |
| Input voltage range (differential) | ±10 V ±30 V | linear range maximum, outside of linear range: max. nonlinearity uncertainty: 300 ns |
| Common mode input voltage | max. ±30 V |  |
| Analog bandwidth | 500 kHz | -3 dB (full power) |
| Analog filter | bypass (without filter), 20 kHz, 2 kHz, 200 Hz | adjustable (globally for all channels) Butterworth, 2nd order |
| CMRR | 70 dB (typ.), 50 dB (min.) 60 dB (typ.), 50 dB (min.) | DC, 50 Hz 10 kHz |
| Gain uncertainty | <1% | 23°C |
| Offset | <1% | 23°C |
| frequency stability | <100 ppm | deterioration < 5 ppm / year |
| Overvoltage protection | ±60 V | long-term |
| Sensor supply | +5 V, 200 mA, 20 mA | Reference: GND |
| Supply voltages | 10 V to 50 V DC |  |
| Sensor supply voltage at CAN/x-INC4-xx-SUPPLY | +2.5 V to +24 V |  |
| Power consumption | 4 W (typ.) | 12 V supply, 23°C |
| Operating temperature | -40°C to 85°C |  |
| Dimensions (W x H x D) | 35 x 111 x 90 mm |  |
| Weight | 300 g |  |
| Terminal connection Rear side | 2x DSUB-15 4x ITT VEAM 2x DSUB-9 2x DSUB-9 PHOENIX (MC 1,5/4STF-3,81) |  |

DSUB Connector for INC4

For customers sensors

Note:

Customer has to iron his sensors/cable glands himself to the connectors

Wall bracket 2.5m for boom-box	1000075710

for the installation of a boom-box incl. installation accessories

Figure 1: drawing of wall bracket

Figure 2: connecting dimensions for installation with steel plate

Figure 3: connecting dimensions for installation without steel plate

| Scope of Supply | Quantity |
| --- | --- |
| Wall bracket 2.5m for boom-box 19" consisting of - flange coupling - tube section 0.5m - angular piece - tube section 2.5m - wall hinge - adapter flange | 1 |
| Installation set consisting of - 6x hexagon screw + nut - 16x washer - 1x steel plate wall support - 4x shear connector for pure concrete wall installation - 1x ground strap 10mm² / 4m for potential equalization | 1 |

| Technical Data |  |
| --- | --- |
| Color | square-type tube: zinc-plated components: RAL 7043 |
| Material | square-type tube: steel 5mm components: AlSi and resp. cast iron |
| Dimensions | see drawing |
| Max. load | see diagram |
| Overall weight [kg] | see drawing |
| Protection class acc. to EN 60529 | IP 65 |

| Interface |  |
| --- | --- |
| Cable routing | inside tube |
| Free space for cable laying [mm] | 70 x 70 |
| Installation with steel plate: - Installation type - Connecting dimensions - Installation material | 4x through boring at wall see drawing with delivered threaded bolt |
| Installation without steel plate: (pure concrete wall) - Installation type - Connecting dimensions - Installation material - Nominal diameter of drill [mm] - Borehole depth [mm] | 4x borehole for screw anchor see drawing with delivered sheer connectors M16x 190 18 125 |

Boom Box 7U	1000012386

for installation of max. 6x 1U 19” racks

Typically for wall mounting in combination with a wall bracket.

| Figure 1: boom-box 7U | Figure 2: application example: photo of an equipped boom box 7U |
| --- | --- |

| Scope of Supply | Quantity |
| --- | --- |
| Boom Box 7U complete mechanical assembled | 1 |
| Slide bar set for 1U | 6 |
| 19” frame (front side) | 1 |
| Hole for cable entry and wall bracket at top | 1 |
| Ground strap for EMC potential compensation | 1 |

| Technical Data |  |
| --- | --- |
| Max. installation depth of racks [mm] | 300 |
| Color | RAL 7035 light grey (front), decor/frame color : RAL 7043 / RAL 7038 |
| Material | AlMgSi corner caps: AlSi |
| Dimension boom box Height [mm] Width [mm] Depth [mm] | 422 535 435 |
| Dimension 19” installation frame | 7U according to DIN 41494-7 (IEC 297-2) for installation of 19”-panels and racks 1U taken for design panel (space for circulation of air) |
| Overall weight [kg] | 40 |
| Protection class housing acc. to EN 60529 | IP 41 (equipped and with installed bracket support) |
| length of ground strap [m] | 4 |

| Interface |  |
| --- | --- |
| Cable entry point | top |

Note:

- Installation space of not ordered 19” racks will be covered with dummy plates.

Cable Set Measuring Module Type D/S 0.3m	1000013434

for connection measuring modules type D/S to another measuring module with max. 1U space between

| Scope of Supply | Quantity |
| --- | --- |
| Supply connecting cable DC 24V | 1 |
| CANopen bus connecting cable | 1 |

| Technical Data |  |
| --- | --- |
| Connector type power cable | D-Sub 9p male/female |
| Connector type bus cable | D-Sub 9p male/female |
| Cable length [m] | 0.3 |

Customized cable length for Measuring Box / Boom Box	1000013436

| Scope of Supply | Quantity |
| --- | --- |
| connection the measuring or boom box to test stand automation system (STARS-PC): Power supply connecting cable DC , CANopen bus connecting cable | 1 |
| CAN bus adaptor with terminating resistor 120Ω | 1 |

| Technical Data |  |
| --- | --- |
| Cable length [m] | [length on request, max 40 m] |

Analog Data Acquisition module LAM002	1000051746

for measuring, conditioning, converting analogue signals of Resistance Thermometer Pt100 or Pt1000, thermocouples NiCr-Ni, voltages or currents. With CAN-Bus interface for connection to STARS test stand automation. Each input can be configured for measuring range and sensor by software.

Figure 1: analogue data acquisition module LAM002

Figure 2: analogue data acquisition module LAM002 dimensions

| Scope of Supply | Quantity |
| --- | --- |
| CAN module LAM 002 | 1 |
| Universal analogue inputs, each with: Amplifiers with differential input Sensor feeding for Pt100 / Pt1000 and potentiometer Internal comparison point for thermocouples 32 point linearization Antialiasing filter coupled with measuring frequency | 8 |
| Analogue / digital converter 16 bit | 1 |
| Controller with CAN bus interface | 1 |
| Diagnostics LEDs | 7 |
| Documentation operating manual | 1 |

| Technical Data |  |
| --- | --- |
| Measuring range thermocouple type K / J [°C] | -20..+1000 |
| Measuring range Pt100 / Pt1000 [°C] | -20..+200 |
| Voltage [V] | 0.1 / 0.2 / 0.5 / 1 / 2 / 5 / 10 |
| Current [mA] | 0..20 / 4..20 |
| Potentiometer [%] | 0..100 |
| Measuring frequency adjustable per channel [Hz] | 1..100 |
| Operating voltage [V] | +24 |
| Operating voltage deviation [V] | +19..+34 |
| Dimensions | Please refer to figure 2 |
| Weight [g] | 790 |

| Interface |  |
| --- | --- |
| Installation type | top hat rail TS 35 |
| Protection class housing acc. to EN 60529 | IP 20 |
| Bus type | CAN open 2.0 |
| Serial service interface | RS232 |
| Electrical connection CAN interface analog inputs power supply | terminal part with screw connection |

Cable Set for Analogue Data Acquisition Module LAM002	1000012213

| Scope of Supply | Quantity |
| --- | --- |
| Connection cable DC 24V and CANopen bus cable to the test stand automation | 1 |

| Technical Data |  |
| --- | --- |
| Cable length [m] | 20 |

Built-in Housing for LAM002	1000007032

for installation of one analogue data acquisition module LAM002 or other similar top hat rail modules in heavy duty environment

Typically for installation close to the specimen on a mounting support

Figure 1: built-in housing for LAM

Figure 2: dimensions of built-in housing for LAM

| Scope of Supply | Quantity |
| --- | --- |
| LAM box housing for 1 module | 1 |
| Screw cover | 1 |
| EMC cable glands | 3 (default) |
| Hole plugs (brass) | 9 (default) |
| Top hat rail (DIN) | 1 |
| Mounting bracket (welded on) | 2 |

| Technical Data |  |
| --- | --- |
| Color | RAL 7035 light grey, structure lacquer |
| Material | sheet steel |
| Sealing (cover) | Foam gasket |
| Dimension of LAM box housing | Please refer to figure 2 |
| Max. installation height [mm] of top hat rail module | 120 |
| Dimension of mounting bracket Length [mm] Width [mm] | 30 30 |
| Dimension of slotted hole [mm] | 15 x 7 |
| Protection class housing acc. to EN 60529 NF C 20-010 Nema | IP 66 IP 669 4, 12, 13 |
| Approval | UL, CSA, GL |

| Interface |  |
| --- | --- |
| Installation system | top hat rail (DIN) |
| EMC gland size for cable diameter [mm] | M16 x 1,5 5..9 (default) |

Weather Station with 15m cable	1000012506

for measuring the environmental conditions (absolute pressure, relative humidity and ambient temperature) of the test cell.

Figure 1: Humidity / temperature and absolute barometric pressure measuring unit

Figure 2: dimensions of weather station

Figure 3: dimensions of pressure sensor

| Scope of Supply | Quantity |
| --- | --- |
| Absolute barometric pressure measuring unit consisting of pressure transducer connector wall-mounting support | 1 |
| Humidity / temperature measuring unit consisting of electronic module combined sensor for relative humidity and air temperature in plastics housing, removable with cable (2m) mounting set | 1 |
| Cable set consisting of measuring cable humidity / temperature measuring unit measuring cable pressure measuring unit supply cable | 1 |
| Documentation: Operating manual Factory calibration certificate | 1x CD, 1x paper |

| Technical Data |  |
| --- | --- |
| Absolute barometric pressure measuring unit Measuring range [kPa] Accuracy (according to IEC 60770) Long term stability at reference conditions Thermal effects (Offset and Span) Compensated temperature range [°C] Cable type measuring cable | +70...+120 ≤ ± 0,05 % FSO BFSL ≤ ± 0,1 % FSO / Jahr ≤ ± 0,05 % FSO / 10K -10 ... 50 LIYCY 2x0.5 mm², shielded wire color code: DIN 47100 |
| Humidity / temperature measuring unit Humidity measuring range [%rH] Measuring accuracy humidity measuring 0...90%rH [%rH] at 20°C Measuring accuracy humidity measuring 90...100%rH [%rH] at 20°C Temperature measuring range sensor [°C] Operating temperature range housing [°C] Measuring accuracy temperature measurement [°C] Cable type measuring cable | 0…100 1 1.7 -20…+80 -20…+60 0.2 LIYCY (TP) 3x2x0.25 mm², shielded wire color code: DIN 47100 |
| Measuring signal pressure [mA] | 4..20 |
| Measuring signal humidity [mA] | 4..20 |
| Measuring signal temperature [mA] | 4..20 |
| Dimension of humidity/temp. measuring unit | Please refer to Figure 2 |
| Dimension of pressure measuring unit incl. wall-mounting support Height [mm] Width [mm] Depth [mm] | 115 90 40 |
| power supply humidity / temperature measuring unit [V DC] | 24 |
| Cable length power and measuring cable [m] | 15 |
| Cable length between humidity and temperature sensor and measuring unit [m] | 2 |
| Cable type measuring cable | LIYCY (TP) 3x2x0.25 mm², shielded wire color code: DIN 47100 |
| Cable type power cable | LIYCY 2 x 1,5 mm², shielded wire color code: DIN 47100 |

| Interface |  |
| --- | --- |
| Type of installation | wall-mounted |
| Pressure connection | ¼“ inside thread |

Customized I/O

-Power supply for Specimen

-I/O for Specimen

-Power Supply for Sensors and Actuators of Facility

-I/O for Sensors and Actuators of Facility

Abb.: Example of I/O Box in Test Cell

Abb.: Example I/O Panel at Operators Place

|  | No. Of Channels | Where? | Connector Type |
| --- | --- | --- | --- |
| AO +/-10V | 4 | I/O Box | BNC |
| Frequency counter, customer wish is TTL as well as up to 24 V DC. 100 kHz | 4 | I/O Box | BNC |
| 8 TTL 5V out | 8 | I/O Box | BNC |
| 8 TTL 5V in | 8 | I/O Box | BNC |
| INC 4, 1 Module | 4 | I/O Box | BNC |
| Netzteil/power supply 5V 10A | 5 pairs | I/O Box | Banana |
| Netzteil/power supply 12V 50A | 5 pairs | I/O Box | Banana |
| Netzteil/power supply 24V 20A | 5 pairs | I/O Box | Banana |
| Ethernet Channels | 4 in front panel | I/O Box | RJ45 |
| Ethernet Channels | 4 in front panel | TITAN Control Cabinet | RJ45 |
| AO +/-10V | 2 | TITAN Control Cabinet | BNC |
| AO 4…20mA | 2 | TITAN Control Cabinet | BNC |
| RS232 serial interface | 1 in front panel, wired to STARS HW | TITAN Control Cabinet | DSUB 9pin |
| RS232 serial interface | 1 in front panel, wired to STARS HW | I/O Box | DSUB 9pin |
| RS232 serial interface | 6 fix on STARS HW | Backside STARS HW | DSUB 9pin |
| CAN interface | 2 in front panel | I/O Box |  |

Automation System

STARS Engine Workstation	1300043538

Workstation Hardware for STARS Automation Software.

Figure 1: STARS Workstation system overview

Figure 2: STARS Workstation (example picture)

Benefits:

The STARS Workstation is specially developed, tested and released for the use with STARS Automation Software and maximizes the performance and robustness of the software.

| Scope of Supply | Quantity |
| --- | --- |
| Process computer with Intel® Xeon® (Embedded) E3-1275 v3 (8MB Cache, 3.5GHz) or better 19” industrial rack 3 x PCI, 3 x PCIe X4, 5 x PCIe X16 slots Solid State Drive (SSD) > 240 GB Onboard Graphics Intel® HD supporting dual monitor output 8 GB RAM | 1 |
| Real-time Ethernet Interface to SPARC-controller | 1 |
| Real-time Ethernet Switch for connection to SPARC Controller | 1 |
| 2-fold CAN-card PCIe | 1 |
| 2x RS232 serial interfaces on PCI board | 1 |
| 24" TFT monitor | 2 |
| Connecting cable for monitor 3m | 2 |
| Optical wheel mouse | 1 |
| Keyboard (US English revision) | 1 |
| Extension cable for key board or mouse 5m | 2 |
| Operating system MS W7 Prof (Multilanguage) | 1 |
| MS Office 2016 Standard, English | 1 |

| Technical Data |  |
| --- | --- |
| Color Front Plate | PB007H light grey |
| Dimension | ATX form factor, 4 HU / 19” acc. to DIN 41494-1 dimensions for 19” racks |
| Protection class housing acc. to EN 60529 | IP 41 (front) |

| Interface |  |
| --- | --- |
| Installation depth [mm] < (without connected plugs) | 450 |
| LAN / Ethernet | 2x onboard 10/100/1000 Mbit 1 x real-time via PCIe card |
| CAN Bus | 2x CAN bus interface via PCIe card (for measuring devices) |
| Serial | 2x serial RS232 via PCI card (for measuring / conditioning devices) |
| USB 2.0 | 3 |
| USB 3.0 | 4 |

Note:

The scope of supply may change depending on the technical development

STARS basic test stand I/O	1000004904

Peripheral test stand I/O of STARS Test Stand Automation System for communication with test stand systems or infrastructure.

Typical for installation at the back panel of a measuring- and control cabinet

Figure 1: Drawing basic test stand I/O

| Scope of Supply | Quantity |
| --- | --- |
| STARS basic test stand I/O 16DI/16DO/2AO | 1 *) |

| Technical Data |  |
| --- | --- |
| DI reserved by STARS for internal test stand control | 4 |
| DO reserved by STARS for internal test stand control | 8 |
| Power supply BUS (electronics) - supply voltage [V DC] - max. supply voltage deviation [%] - max. current consumption [mA] | 24 -25..+30 500 |
| Power supply field - supply voltage [V DC] - max. supply voltage deviation [%] - max. current feeding [A] | 24 -25..+30 10 |

| Interface |  |
| --- | --- |
| Free DI 24V DC, pos. logic, two- or three-wire system | 8 |
| Free DO 24V DC, pos. logic, max. 0.5 A short-circuit proof | 8 |
| Free digital limit contact inputs (DI predefined by STARS) 24V DC, pos. logic, two- or three-wire system | 4 |
| Free analog outputs (± 10V DC) | 2 |
| Connection type | spring clamps |
| Wire cross section [mm² (AWG)] | 0.08..2.5 (28..14) |
| Bus type | CANopen |

Note:

If the measuring and control Cabinet type S or T/D is purchased, this module will be delivered installed in the cabinet

Direct wiring of the 24V DC supply of the test stand I/O is recommended, if the module is installed inside another cabinet.

STARS Engine Workstation License	1000012410

The STARS Engine domain combines the STARS platform with specific functions and structures suitable for engine testing.

Gathering the real time function and user interface together in one integrated system environment greatly enhances the interaction and thus the productivity and quality of testing related activities.

STARS is a high performance, flexible automation platform with integrated automation, data acquisition and control functions that provides a broad range of application functionality in a single, powerful environment.

Figure 1: STARS GUI

Due to its in-build flexibility STARS Engine is especially suited for a wide range of engine testing applications supporting the whole development process from component development to certification. The range of possible applications includes:

Durability Testing

Performance and Emissions Development

Quality Audit

Mechanical Development

Fuel and Lubricants Testing

Noise, Vibration and Harshness

Heavy Duty Engine Emissions

Light Duty Engine Emissions

STARS Automation provides:

-	Defining and executing user-level tests (called Workflows in STARS) including preparation, complete test activity, data flow processes, report generation and other customer specific actions

-	Test preparation and execution of manually or automatically operated tests.

-	SPARC Engine controller is fully integrated into STARS Engine. Visualization, remote control, configuration and parameterization of the controller is supported via STARS

-	Safe control and monitoring of the specimen in all modes of operation

- Control of test stand and test stand devices

-	Data acquisition including integration of special measuring systems and intelligent instruments for engine testing, such as Fuel Meter, Combustion Analyzer, Emission Analyzer and more

- Communication with TCU via application systems like INCA or direct communication with the vehicle bus system via CAN or other bus systems

- Acquisition, monitoring, visualization and reporting of measured or derived parameter, such as speed, torque, power, fuel consumption, emissions, etc.

-	Logging of events during the test, such as system messages and user entered comments.

KEY ADVANTAGES

Powerful

Execution of test schedules, monitoring, scripts in real-time with up to 1 kHz due to the in-build STARS real-time system

Data acquisition with high throughput of more than 150.000 Values/ sec and high sampling rates up to 5 kHz / channel

Future proof due to in-build flexibility

No license restrictions on the number of standard names, channels, data logger, monitoring channels, calculations, scripts etc. that can execute concurrently

Project and test stand structure can be setup to the customer requirements

STARS Resources can be used in STARS projects and in STARS test stands

Several instances of STARS resources e.g. test schedules or limit data sets can be used at the same time

Open

Open software: The STARS User can write scripts which can be executed in STARS in real time with up to 1 kHz. STARS includes a generic interface to communicate with external software via Microsoft Visual Studio Tools for Applications.

Open interfaces: Support of interface standards like CANopen and EtherCat for I/O connection, ASAM ODS Interface for evaluation tools

Due to the market presence of STARS many interfaces to commonly used devices already exist.  In addition to the steadily growing number of interfaces STARS Users have the possibility to develop customized interfaces with the optional driver development framework.

Global Support

The STARS automation platform is a HORIBA core product and will continuously be extended and updated.

Continuity: It is HORIBA philosophy to be backwards compatible. Test schedules developed for the first STARS Version in 2003 can still be executed on actual STARS versions.

HORIBA has an existing global network of STARS support centers to deliver and support STARS systems in all relevant automotive markets

The system is delivered pre-configured. Its configuration is related to the delivered hardware components and to the pre-defined interfaces.

The software includes default tests and a basic setup to control engine test stands in manual and automatic mode.

In a manual test, operating points can be entered in the STARS panel and measurements can be triggered manually by clicking a button.

Figure 2: Manual Measurement

This allows to execute steady state tests in manual mode and to kick-off a defined measurement under stable conditions. Continuous signals and device communication are coordinated. The test consists of these elements:

Workflow to load all necessary components

Panel for user interaction with the test stand enabling the following functions

Start/Stop of the engine

Ignition On/Off

Selection of control modes

Setting demand values for the engine parameters

Start measurement

Live display of actual engine operation data

Test programs for the measurement procedure

Data logging with user definable channel list

Excel based Report to present all measured values

In the standard configuration an example of an automated test procedure with user guidance is included. The test includes the following steps:

Input of test parameters

Panel for presentation of live data and for user interaction

Automatic start and stop of the engine

Warm up engine with steps depending on user defined cooling water temperature

Power map test

Engine cool-down

Excel based report of the test result

| Scope of Supply | Quantity |
| --- | --- |
| STARS Engine Workstation License | 1 |
| Software license for real-time extension RTX | 1 |
| STARS Device driver for test stand controller SPARC Engine | 1 |
| STARS Device driver for CAN based I/O (Analogue and digital In- and output) | 1 |
| STARS help desk for the warranty period (max.1 year) | 1 |

| Technical Data |  |
| --- | --- |
| Functions | See Technical Specification STARS |
| Operating System | Microsoft Windows 7 32-bit |

STARS Device Driver License Opacimeter AVL 439	1000007055

Software Driver for interfacing the Opacimeter AVL 439. The device driver will allow remote control of the Instruments out of any STARS application. The STARS Device Driver plug-in concept will allow using this device without additional programming. Once the device driver is integrated into the test stand configuration it could be used from any STARS test schedule or display.

| Scope of Supply | Quantity |
| --- | --- |
| Driver for Opacimeter AVL 439 | 1 |
| Manual Control Display Page | 1 |
| Embedded User Guide | 1 (US-English only) |

| Technical Data |  |
| --- | --- |
| Functions | Support commands to control the system and to transfer measured values and status information |
| Driver/Instrument Class | Smoke Measurement |
| Devices supported | AVL 439 |
| Device Driver Capabilities | The following values are updated @ 1Hz: Device status Interface mode Active alarms Gas Temperature Measuring Chamber Pressure Gas Flow Rate Measuring Chamber Temperature Detector Temperature Detector Signal Halogen Lamp 1 & 2 Current Lamp Temperature Heat Output Window 1 & 2 Warmup completion time The following values are updated @10Hz: Opacity Absorption The following values are updated at device connection and when reconfigured by the driver: Variable to output Variable to filter Filter type Rise time Trigger type The following values are updated at device connection: Firmware version Device ID The following values are updated at device connection and every 5 minutes: Operating time Time since service |

| Interface |  |
| --- | --- |
| Platform | STARS Version 1.6 onwards |
| Protocol | AK |
| Physical Interface | Serial RS-232 |

| Test Schedule Plug-In Commands | Description |
| --- | --- |
| Connect | Connect device to STARS |
| Release | Disconnect device from STARS |
| SetMode | Sets the mode of the instrument |
| Measure | Starts a continuous measurement |
| StartPeakMeasurement | Starts peak measurements |
| EndPeakMeasurement | Stops peak measurements |
| PauseInstrument | Suspends operation of the instrument |
| Off | Stops all functions and turns off heaters |
| Zero | Initiates a warm-up phase |
| CheckZero | Flows clean air through the instrument |
| CheckLinearity | Performs a linearity check of the instrument |
| Purge | Performs a back flush of the probe line and sampling line |

STARS Device Driver Software License Fuel Flow Meter AVL	1000007061

Software Driver for Interfacing the Fuel Flow Meter AVL 733 and 733S. The device driver will allow remote control of the Instruments out of any STARS application. The STARS Device Driver plug-in concept will allow using this device without additional programming. Once the device driver is integrated into the test stand configuration it could be used from any STARS test schedule or display.

| Scope of Supply | Quantity |
| --- | --- |
| License for: Fehler! Unbekannter Name für Dokument-Eigenschaft. Device Driver | 1 |
| Manual Control Panel | 2 |
| Embedded User Guide | 1 (US-English only) |
| STARS Help desk for the warranty period (max.1 year) | 1 |

| Technical Data |  |
| --- | --- |
| Driver/Instrument Class | Fuel Measurement |
| Devices Supported | AVL 733 AVL 733S |
| Device Driver Capabilities | MassConsumed: Mass of fuel consumed for the last completed measurement @ 2 Hz MeasurementTime: Time interval of the last completed measurement @ 2 Hz InstMassFlowRate: Instantaneous mass flow rate @ 5 Hz MassFlowRate: Mean mass flow rate over the last completed measurement @ 2 Hz AverageMassFlowRate: Mean mass flow rate across a number of measurements @ 2 Hz SDOfMassFlowRate: Standard deviation of mass flow rate across a number of measurements @ 2 Hz InstantaneousMass: InstantaneousMass @ 2 Hz MeasurementNumber: Number of the current measurement @ 2 Hz CurrentTime: Current time setting for time-controlled measurements @ 2 Hz |

| Interface |  |
| --- | --- |
| Platform | STARS Version 1.6 onwards |
| Protocol | AK |
| Physical Interface | Serial RS-232 or Ethernet |

| Test Schedule Plug-In Commands | Description |
| --- | --- |
| Set Mode | Set Mode: Mass, Time, Continuous or Manual |
| Measure | Start a continuous measurement |
| Start | Start a manual measurement |
| Stop | Stop manual measurement |

STARS Device Driver AVL Indicom	1000074911

Software Driver for interfacing the combustion analyser AVL Indicom via DCOM interface. The device driver will allow to remote control the Indicom and to transfer measured data to STARS.

The STARS Device Driver plug-in concept will allow using this device without additional programming. Once the device driver is integrated into the test stand configuration it could be used from any STARS test schedule or display

The device driver supports Indicom from version 1.2 and later with options R and X[D] installed.

| Scope of Supply | Quantity |
| --- | --- |
| License for: AVL Indicon Device Driver | 1 |
| Manual Control Panel | 1 |
| Embedded User Guide | 1 (US-English only) |
| STARS Help desk for the warranty period (max.1 year) | 1 |

| Technical Data |  |
| --- | --- |
| Functions | Support the parameterisation and remote control of the device and the transfer of measured values and status information |
| Driver/Instrument Class | External Data Logging |

| Device Driver Capabilities |  |
| --- | --- |
| Data Rate | Update @ 1 Hz |
| Measured values | Maximum of 10 Hz (configurable) |
| Active alarms | By callback |
| Device status | By callback |
| Firmware version | Updated at device connection |
| Data Rate | Update @ 1 Hz |

| Interface |  |
| --- | --- |
| Platform | STARS Version 1.8 onwards |
| Protocol | Microsoft DCOM |
| Physical Interface | Ethernet |

| ULI Methods - Commands | Description |
| --- | --- |
| Connect | Connect device to STARS. |
| Disconnect | Disconnect device from STARS. |
| Remote | Switch device into remote controlled mode. |
| Local | Switch device into locally controlled mode. |
| StartMonitor | Start continuous measurement. |
| StopMonitor | End continuous measurement. |
| SetBatchSize | Set number of cycles for batch measurement. |
| AcquireBatch | Run batch measurement. |
| AcquireAndSaveBatch | Run batch measurement and save results. |
| CancelMeasurement | Stop any running measurements. |
| SaveData | Save measured data. |
| LoadConfigFile | Load device configuration from file. |
| GetErrorString | Convert ULI error number into text. |

STARS Device Driver ASAM ASAP3	1000007064

Software Driver for Interfacing ECU application systems via ASAP3 allows to measure data and change scalar or map data to/from an ECU application system.

The device driver supports the following devices:

ETAS INCA (4.0.x, 5.1 to 5.x)

ATI Vision (2.1 to 2.5)

Gredi

Vector CanAPE (5.x and 6.x)

Denso Panel 3

The STARS Device Driver plug-in concept will allow using this device without additional programming. Once the device driver is integrated into the test stand configuration it could be used from any STARS test schedule or display.

Figure 1: Schematic diagram STARS-interface ASAP3

| Scope of Supply | Quantity |
| --- | --- |
| STARS Driver ASAP3 interface | 1 |
| Display screen for visualisation the signals exchanged via ASAP3 | 1 |
| Application support | Not included, will be quoted separately |
| STARS Help desk for the warranty period (max.1 year) | 1 |

| Technical Data |  |
| --- | --- |
| Functions | Send and receive variables to/from ASAP3 Devices |
| Access speed for variables [Hz]* | 10; 5; 2; 1; 0,5; 0,2 |
| Access speed for parameters [Hz]* | 10; 5; 2; 1; 0,5; 0,2 |
| Max. number variables and parameters to be supported | 400 |

| Interface |  |
| --- | --- |
| Platform | STARS Automation |
| Protocol | ASAP3 2.1 (without extended functions) |
| Physical Interface | Serial RS232 or TCP/IP |

| Acceptance Test |  |
| --- | --- |
| Establish Connection to Application system | Setup connection to customer’s application system. If such system is not available, a connection to an INCA system running in Demo mode on the PC of the commissioning engineer will be established |
| Setup a limited number (max. 20) channels for measurement from application system, show online values in STARS variables in a display page. Control of driver is done by use of default display page for the driver | If such system is not available, a setup of two channels from INCA in Demo mode for proof of capability will be performed |

* depending on the amount of data and the performance of the ASAP3-device as well as the physical interface and the load of the STARS system.

Installation STARS Device Driver HORIBA MEXA ONE LAN	1000007058

Installation of a software driver for interfacing the HORIBA MEXA ONE series. The device driver will allow remote control of the Instruments out of any STARS application. The STARS Device Driver plug-in concept will allow using this device without additional programming. Once the device driver is integrated into the test stand configuration it could be used from any STARS test schedule or display.

| Scope of Supply | Quantity |
| --- | --- |
| Software license and installation of Driver for HORIBA MEXA ONE LAN | 1 |
| Setup display screen for manual interaction with the device | 1 |

| Technical Data |  |
| --- | --- |
| Functions | Support of the remote control commands of the HORIBA MEXA ONE series and to transfer measured values and status information |

| Interface |  |
| --- | --- |
| Platform | STARS Automation |
| Protocol | MEXA LAN |
| Physical Interface | Ethernet |

Note:

The scope of supply consist of the software license (free of charge) and a handling charge

Installation STARS Device Driver HORIBA MEXA 1000SPCS	1000044868

Installation of a software driver for interfacing the HORIBA MEXA 1000SPCS series. The device driver will allow remote control of the Instruments out of any STARS application. The STARS Device Driver plug-in concept will allow using this device without additional programming. Once the device driver is integrated into the test stand configuration it could be used from any STARS test schedule or display.

| Scope of Supply | Quantity |
| --- | --- |
| Software license and installation of Driver for HORIBA MEXA 1000SPCS | 1 |
| Display screen for manual interaction with the device | 1 |

| Technical Data |  |
| --- | --- |
| Functions | Support of the commands to remote control the HORIBA MEXA 1000SPCS series and to transfer measured values and status information |

| Interface |  |
| --- | --- |
| Platform | STARS Automation |
| Protocol | AK |
| Physical Interface | Ethernet |

Note:

- The scope of supply consist of the software license (free of charge) and a handling charge

Installation STARS Device Driver HORIBA MDLT 1300T	1000007062

Installation of a software driver for interfacing the particulate-emission measurement system HORIBA MDLT 1300T series. The device driver will allow remote control of the Instruments out of any STARS application. The STARS Device Driver plug-in concept will allow using this device without additional programming. Once the device driver is integrated into the test stand configuration it could be used from any STARS test schedule or display.

The Horiba MDLT STARS Device Driver supports the remote control commands of the models:

MDLT-1302T

MDLT-1303T

| Scope of Supply | Quantity |
| --- | --- |
| Software license and installation of Driver for HORIBA MDLT 130xT | 1 |
| Editor for device configuration and Display screen for manual interaction with the device | 1 |

| Technical Data |  |
| --- | --- |
| Functions | Start / Stop of the blower and the supply of dilution air Control of the particulate sampling Monitor of Instantaneous data Acquisition and reset of the integrated data Setting dilution air flow control mode and the condition Control of the auto filter changer Start / Stop of the learning of exhaust flow input signal Setting of the Input signal |

| Interface |  |
| --- | --- |
| Platform | STARS Automation |
| Protocol | HORIBA MetaLAN |
| Physical Interface | Ethernet |

Note:

The scope of supply consist of the software license (free of charge) and a handling charge

Due to the concept of particular measurement with HORIBA MDLT the interface is used to control the device but not evaluate the particulate emission results. This is typically done with the user software of the HORIBA MDLT.

Installation STARS Device Driver HORIBA PTFM 1000	1000059149

Installation of a software driver for interfacing the HORIBA PTFM 1000 series exhaust mass flow meter. The device driver will allow remote control of the Instruments out of any STARS application. The STARS Device Driver plug-in concept will allow using of this device without additional programming. Once the device driver is integrated into the test stand configuration it could be used from any STARS test schedule or display.

| Scope of Supply | Quantity |
| --- | --- |
| Software license and installation of Driver for HORIBA PTFM 1000 | 1 |
| Display screen for manual interaction with the device | 1 |

| Technical Data |  |
| --- | --- |
| Functions | Support of the full set of remote control commands of the HORIBA PTFM 1000 series and to transfer of measured values and status information |

| Interface |  |
| --- | --- |
| Platform | STARS Automation |
| Protocol | HORIBA |
| Physical Interface | Ethernet |

- The scope of supply consist of the software license (free of charge) and a handling charge

Device Driver Software License Blow-By Meter 442 for STARS Test Stand Automation	1000007044

Software Driver for Interfacing the Blow-By Meter 442. The device driver will allow remote control of the Instruments out of any STARS application. The STARS Device Driver plug-in concept will allow using this device without additional programming. Once the device driver is integrated into the test stand configuration it could be used from any STARS test schedule or display.

| Scope of Supply | Quantity |
| --- | --- |
| Driver for Blow-By Meter 442 | 1 |
| Display screen and parameterisation for manual interaction with the device | 1 |
| - Technical Data |  |
| Functions | Support of commands to control the system and to transfer measured values and status information |
| Interface |  |
| Platform | STARS Automation |
| Protocol | AK |
| Physical Interface | Serial RS-232 |

STARS RS 232 Serial  Board	1300033783

STARS Interface for up to 8 serial connections

| Scope of Supply | Quantity |
| --- | --- |
| Serial Connection Board | 1 |
| Connection Cable 1 m | 1 |
| Technical Data |  |
| Number of Serial connections | 8 |
| Protocol | RS 232 |
| Max. Baudrate | 921.6 kbps |
| Interface |  |
| Connector | 8 x DB9 male |

STARS High Speed CAN Interface	1300033780

Connection board for communication between STARS and up to 2 CAN segments

| Scope of Supply | Quantity |
| --- | --- |
| CAN Connection Board | 1 |
| Technical Data |  |
| Number of CAN connections | 4 |
| CAN protocol | ISO 11898-2 |
| Interface |  |
| Connector | 9-pin DSUB male |

Operators Place

Measuring and Control Cabinet Type T/D	1300042242

for installation of 19” rack modules and test stand electrics

Complete assembled cabinet for engine test stands with AC machine (e.g. Dynas3 loading unit), if purchased together with the corresponding module. Installation space of not ordered 19” racks will be covered with dummy plates.

Figure 1: measuring and control cabinet type T/D fully equipped 
(design similar to illustration)

Abb. 2: dimensions of movable cabinet 19” / 31HU

| Scope of Supply | Quantity |
| --- | --- |
| Moveable Cabinet 19” / 31HU consisting of: mech. complete assembled 4 x cabinet roller 19” frame (front sided) subplate (back sided), swiveling hinge door (back sided) 4 x eye bolt, zinced 5 x slide bar set 8 x flat strip earthing shield bus cable feed from below / back | 1 |
| Internal cable set + wiring with CE cables | 1 |
| Power supply consisting of 1 19” main switch front panel 1 1x 20A mains filter 1 multiple socket outlet CEE 7-4 1 AC terminal strip 1 DC terminal strip 2 24V DC power supply unit 10A 4 automatic circuit breakers | 1 |
| Power contactor 3 N.C.C. (forcibly guided contacts ) for remote power disconnection e.g. for fuel consumption measuring device | 1 |
| Installation space incl. cabling for STARS test stand I/O air flow meter test stand controller type SPARCE Central Safety Module SZM control and power unit throttle position type LSR2003 | 2 1 1 1 1 |
| Documentation according to EG-machine-directives and the resulting standards DIN EN 61000-6-2, DIN EN 55011, DIN EN 60204-1 | 1 |
| Safety test report acc. to BGV A3 / DIN EN 60204-1 | 1 |

| Technical Data |  |
| --- | --- |
| Single wire colors acc. to DIN EN 60204-1 (IEC 60204-1) | main circuit: black neutral conductor: light blue main circuit in front of main switch: orange DC control circuits: dark blue AC control circuits: red protective conductor: green/yellow |
| Circuit fusing | acc. to DIN EN 60204-1 |
| Wire cross section | acc. to DIN EN 60204-1 |
| Naming electrical equipment with system, location and identification of equipment | acc. to IEC 61346-1 (1996), DIN EN 6779-1 (1992-09), DIN EN 61346-2 (2000-12), DIN 61346-1 (1997-01) |
| Rated current [A] approx. (full equipped cabinet) | 16 |
| Environmental conditions | +5..+40°C rel. air humidity 5…+85% acc. to DIN EN 60721-3-3 / class 3K3 |
| Terminal strip, type of clamps | spring clamps |
| Dimension of cabinet [mm] | Please refer to figure 2 |
| Dimension 19” installation frame | 31 HU according to DIN 41494-7 (IEC 297-2) |
| Max. installation depth of racks [mm] | 450 |
| Color | grey RAL 7047, powder lacquer |
| Material | sheet steel, |
| Locking mechanism | 4-point |
| Protection class housing acc. to EN 60529 | IP 21 |

| Interface |  |
| --- | --- |
| Free installation space at subplate | approx. 25% |
| Mains type | TN-S |
| Mains connection data | 1/N/PE 230V AC 50 Hz |
| Connected Load S [kVA] | 5 |
| pre-fuse [A] (in combination with Dynas3 installed in VFD cabinet) | 16 |
| Recommended connection line acc. to DIN EN 60204-1 | 3x2.5 mm² (local electric regulations must be taken into consideration) |
| The BUS cable set has a length of 8m, it means maximum clear width between cabinet and operator desk [m] | 3 |
| Cabinet mechanics: Point transmission of force Floor properties Evenness Air outlet B 4.1 Air inlet B 4.2 Cable entry point B 4.3 | over four 65mm dia. rollers mounted in the corners floated concrete max. 2mm/m (gauge) louvers at cabinet top louvers at right cabinet side feed to rear from below. |

Note:

- Installation space of not ordered 19” racks will be covered with dummy plates.

Central Safety Module SZM light	1000112842

safety module (19“) ready for connection with emergency-stop circuit, safety circuit and display functions. Emergency stop button will be connected directly.

Figure 1: central safety module SZM light (design similar to illustration)

| Scope of Supply | Quantity |
| --- | --- |
| Central safety module with LED display field | 1 |
| Documentation operating manual central safety module SZM | 1 |

| Technical Data |  |
| --- | --- |
| Color | RAL 7035 light grey |
| Material front panel | aluminum (AlMg) |
| Dimension measuring rack | 2 HU / 19” acc. to DIN 41494-1 dimensions for 19” racks |
| Installation depth [mm] | 312 |
| Supply voltage DC [V] | 24 |
| Protection class housing acc. to EN 60529 | IP 20 |
| Safety-related characteristic data in acc. with EN ISO 13849-1: 2008 | PL e Cat. 4 |

| Interface |  |
| --- | --- |
| Number of inputs for emergency-stop with positively driven contacts | 5 |
| Number of outputs for emergency stop for the integration of safety-relevant switch-offs (e.g. test stand controller) | 6 |
| Number of inputs for control circuit | 3 |
| Number of outputs for control circuit | 1 |
| Malfunction acknowledgement input for E-Stop´s and safety circuits | 1 |
| Input external malfunction | 1 |

Emergency Stop Button (10m)	1300006412

element for activating the emergency stop circuit of the central safety circuit module SZMlight.

Typical location for 10m version: operator desk

Figure 1: emergency stop button

Figure 2: dimensions of emergency stop button

| Scope of Supply | Quantity |
| --- | --- |
| Emergency stop button with labeling | 1 |
| Electrical contacts: Brake contact Make contact | 2 1 |
| Included connecting cable witch connector for central safety circuit module type SZMlight | 1 |

| Technical Data |  |
| --- | --- |
| Tamper-proof | acc. to ISO 13850 / EN 418, pull to release |
| Auxiliary contacts | safety function by positive opening to IEC/EN 60947-5-1 |
| Language of labeling | English, German, French, Italian |
| Color button / housing | RAL 3000 / yellow |
| Protection class housing acc. to EN 60529 | IP66 |
| Cable length [m] | 10 |

Emergency Stop Button (20m)	1300006417

element for activating the emergency stop circuit of the central safety circuit module SZMlight.

Typical location for 20m version: beside test cell door

Figure 1: emergency stop button

Figure 2: dimensions of emergency stop button

| Scope of Supply | Quantity |
| --- | --- |
| Emergency stop button with labeling | 1 |
| Electrical contacts: Brake contact Make contact | 2 1 |
| Included connecting cable witch connector for central safety circuit module type SZMlight | 1 |

| Technical Data |  |
| --- | --- |
| Tamper-proof | acc. to ISO 13850 / EN 418, pull to release |
| Auxiliary contacts | safety function by positive opening to IEC/EN 60947-5-1 |
| Language of labeling | English, German, French, Italian |
| Color button / housing | RAL 3000 / yellow |
| Protection class housing acc. to EN 60529 | IP66 |
| Cable length [m] | 20 |

Multiple Socket Outlet EMC CEE 7-4 (8m)	1000010550

typically power supply of test stand PC (i.e. STARS-PC) and peripheral equipment at the operator desk for supply from the measuring and control cabinet

Figure 1: multiple socket outlet EMC CEE 7-4

Figure 2: dimensions of multiple socket outlet EMC CEE 7-4

| Scope of Supply | Quantity |
| --- | --- |
| Multiple socket outlet | 1 |
| Connected power cable with free end | 1 |
| Integrated line filter for protection against high frequency power line interference | 1 |

| Technical Data |  |
| --- | --- |
| Cable type | 3x 1,5 mm², unshielded numbered wire code (VDE 0293) |
| Cable jacket color | grey |
| Max. Voltage [V] | 250 |
| Mains frequency [Hz] | 50 / 60 |
| Dimension | Refer to figure 2 |
| Cable length [m] | 8 |
| Filter attenuation 1MHz/100MHz | > 25dB /25dB symm. and asymm. |

| Interface |  |
| --- | --- |
| Socket type | Schuko (CEE 7-4) |
| Number of sockets | 5 |

Power Cable for Measuring and Control Cabinet T/D	1000067281

power supply cable of measuring and control cabinet T/D

Typical for power feeding from inverter cabinet

| Scope of Supply | Quantity |
| --- | --- |
| Power supply cable AC 230V /16A | 1 |

| Technical Data |  |
| --- | --- |
| Cable type power cable | Y-CY-JZ (Oelflex classic 110CY) 3 x 2,5 mm², shielded wire color code: VDE 0290 black wire with white numbers; PE-line green-yellow |
| Cable length power cable [m] | 20 |

Services

Dokumentace papírová i na CDROM	1000010462

figure 1: technical documentation

| Scope of Supply | Quantity |
| --- | --- |
| Engineering services consisting of: - electrical drawings for the complete test stand, integration of the ordered test stand modules - Work out a specification to support the customers on site preparations and technical interfaces of the test stand installation - Preparation of a mechanical assembly drawing of the test stand including the ordered test stand modules - Test stand manual and operating instructions according to the supplied test equipment | 1 |
| Technical documentation according to EC Machinery Directive | 1x CD-ROM, 1x paper |

Instalace na místě

V ceně

Uvedení do provozu

V ceně

STARS  Software-uvedení do provozu na místě

V ceně

| Ausrüstung | Schnittstelle | Protokoll |
| --- | --- | --- |
| Micro Motion Massendurchflusssensor | RS 485 / RS232 |  |
| Abgasemissionsmessung - Horiba Mexa One | RJ45 | AK protokol |
| Abgasdurchflussmessung - Horiba PTFM-1000V2 | RJ45 | AK protokol |
| Opacimeter AVL 439 | RS 232 / RJ45 / 0-10V | AK protokol |
| Horiba MDLT-ONE | RJ45 | AK protokol |
| HORIBA MEXA-2200SPCS | RJ45 | AK protokol |
| Kraftstoffverbrauchsmessung - AVL 733S | RS 232 |  |
| Lambda meter ETAS | RS 232 |  |
| Festphasenspektrometer TSI EEPS 3090 | RS 232 |  |
| DiagRAM Diagnosesoftwareverbindung | RS 232 / USB (vyčítání alespoň 40 kanálů) |  |
| DiagRAM Diagnosesoftwareverbindung | RS 232C |  |
| Indometer (Anzeige von dynamischen Wellenformen) | RS232 |  |

List of devices which will be integrated

Zaškolení min. 8 osob a předání

V ceně